## parameters
parameter | name | longname | description | process | fbnr | vartype | vartype2 | default | min | max | read | write | poll | secured | highly_secured | available | advanced | channels_group0 | channels_group1 | channels_group2 | value_options_count | value_options
1 | identstrng | Identification string | identnr.+softwareversion[+serialnr.] | 0 | 0 | c |  | 7SN999999 |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 13 |  |  | 0
2 | pna | Primary node address | primary node address: network parameter FLOW-BUS | 0 | 1 | c |  | 0 | 0 | 128 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 1 |  |  | 0
3 | sna | Secondary node address | secundary node address: network parameter FLOW-BUS | 0 | 2 | c |  | 0 | 0 | 128 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 1 |  |  | 0
4 | nna | Next node address | next node address: network parameter FLOW-BUS | 0 | 3 | c |  | 1 | 0 | 128 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 1 |  |  | 0
5 | lna | Last node address | last node address: network parameter FLOW-BUS | 0 | 4 | c |  | 32 | 0 | 128 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 1 |  |  | 0
6 | arbitrage | Arbitrage | FLOW-BUS arbitrage setting and/or automatic optimization | 0 | 5 | c |  | 67 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 1 |  |  | 5 | 1: temporary busmaster ¦ 2: always busmaster ¦ 3: automatic busmaster ¦ 67: auto busmaster and auto bus optimalization (fast token ring) ¦ 83: auto busmaster and auto bus optimalization with fail safe
7 | initreset | Initreset | init and reset security key commands for network/parameter settings | 0 | 10 | c |  |  | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 12 |  |  | 0
8 | measure | Measure | measured value (100% = 32000) |  | 0 | i |  | 0 | -23593 | 41942 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 2 |  |  | 0
9 | setpoint | Setpoint | setpoint: wanted value (100% = 32000) |  | 1 | i |  | 0 | 0 | 32767 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 2 | 18 |  | 0
10 | setpslope | Setpoint slope | setpoint ramp signal 0..100 % in up to slope x 0.1 sec. |  | 2 | i |  | 0 | 0 | 30000 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 18 |  |  | 0
11 | analoginp | Analog input | analog input signal, normally used for ext. setp. (100% = 32000) |  | 3 | i |  | 0 | -23593 | 41942 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 2 | 18 |  | 0
12 | cntrlmode | Control mode | control mode selection for instrument or module |  | 4 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 18 |  |  | 25 | 0: setpoint = BUS setpoint ¦ 1: setpoint = analog input ¦ 2: setpoint = master output(FLOW-BUS) * slave factor(FLOW-BUS) ¦ 3: close valve ¦ 4: controller idle (no reaction on changes in sensor signal) ¦ 5: testmode enable (select subject with par 70) ¦ 6: tuningmode enable (select subtject with par 79) ¦ 7: setpoint = 100% ¦ 8: purge valve (fully open) ¦ 9: calibration mode enable (select subject with par 58) ¦ 10: setpoint = master output(analog in) * slave factor(FLOW-BUS) ¦ 11: setpoint = keyboard OR FLOW-BUS setpoint ¦ 12: setpoint = 0% ¦ 13: setpoint = master output(FLOW-BUS) * slave factor(analog in) ¦ 14: (FPP) Range select mode ¦ 15: (FPP) Manual start sensor select, automatic end sensor ¦ 16: (FPP) Automatic start sensor select, manual end sensor ¦ 17: (FPP) Automatic start and end sensor ¦ 18: setpoint = RS232 setpoint ¦ 19: RS232 broadcast mode ¦ 20: valve stearing (valve = setpoint) ¦ 21: analog valve stearing (valve = analog setpoint) ¦ 22: valve safe state ¦ 23: Use fluidset conditions for conversion (instead of p354) ¦ 24: vent DEPx (when available)
13 | polycnst A | Polynomial constant A | polynomial constant A (offset) |  | 5 | f |  | 0.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
14 | polycnst B | Polynomial constant B | polynomial constant B (span] |  | 6 | f |  | 1.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
15 | polycnst C | Polynomial constant C | polynomial constant C |  | 7 | f |  | 0.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
16 | polycnst D | Polynomial constant D | polynomial constant D |  | 8 | f |  | 0.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
17 | polycnst E | Polynomial constant E | polynomial constant E (offset) for setpoint or power value |  | 9 | f |  | 0.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
18 | polycnst F | Polynomial constant F | polynomial constant F (span) for setpoint or power value |  | 10 | f |  | 1.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
19 | polycnst G | Polynomial constant G | polynomial constant G for setpoint or power value |  | 11 | f |  | 0.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
20 | polycnst H | Polynomial constant H | polynomial constant H for setpoint or power value |  | 12 | f |  | 0.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
21 | capacity | Capacity 100% | readout value at 100% in capacity (readout) unit |  | 13 | f |  | 1.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 | 19 |  | 0
22 | sensortype | Sensor type | sensor type information for actual reading and sensor/controller indication |  | 14 | c |  | 3 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 22 | 0: pressure (controller) ¦ 1: liquid volume (controller) ¦ 2: liquid/gas mass (controller) ¦ 3: gas volume (controller) ¦ 4: other sensor type (controller) ¦ 5: temperature (controller) ¦ 6: time (duration) (controller) ¦ 7: frequency (controller) ¦ 8: weight (controller) ¦ 9: solutions (controller ¦ 10: density (controller) ¦ 128: pressure (sensor) ¦ 129: liquid volume (sensor) ¦ 130: liquid/gas mass (sensor) ¦ 131: gas volume (sensor) ¦ 132: other sensor type (sensor) ¦ 133: temperature (sensor) ¦ 134: time (durarion) (sensor) ¦ 135: frequency (sensor) ¦ 136: weight (sensor) ¦ 137: solutions (sensor) ¦ 138: density (sensor)
23 | capunit | Capacity unit index | pointer to capacity (readout) unit (classic unit table] |  | 15 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 | 19 |  | 0
24 | fluidnr | Fluidset index | fluid number: pointer to fluidset (for e.g. polynome, name and cap.) |  | 16 | c |  | 0 | 0 | 8 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
25 | fluidname | Fluid name | name of fluid |  | 17 | c |  | AIR |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
26 | claimnode | Claim node | node address of module with operation rights |  | 18 | c |  | 0 | 0 | 128 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 12 |  |  | 0
27 | modify | Modify | contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed) |  | 19 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
28 | alarminfo | Alarm info | status information of several alarms/errors in the instrument |  | 20 | c |  | 0 | 0 | 255 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 4 |  |  | 16 | 0: no error message in alarm error status register ¦ 0: no warning message in alarm warning status register ¦ 0: no minimum alarm message (measure>minimum limit) ¦ 0: no maximum alarm message (measure<maximum limit) ¦ 0: batch counter has not reached its limit ¦ 0: response O.K. (setpoint-measure within limit) ¦ 0: master output signal O.K. (or not used) ¦ 0: hardware O.K. ¦ 1: at least 1 error message in alarm error status register ¦ 1: at least 1 warning message in alarm warning status register ¦ 1: minimum alarm message for measured signal ¦ 1: maximum alarm message for measured signal ¦ 1: batch counter has reached its limit ¦ 1: response alarm message: setpoint-measure is too high ¦ 1: master output signal not received: check master instrument ¦ 1: hardware error message: check your hardware
29 | chanamount | Channel amount | amount of channels which can be operated | 0 | 12 | c |  | 32 | 1 | 120 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
30 | firstchan | First channel | first channel that can be operated | 0 | 13 | c |  | 1 | 1 | 120 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
31 | lastchan | Last channel | last channel that can be operated | 0 | 14 | c |  | 32 | 1 | 120 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
32 | hostcontrl | <hostcontrl> | operation by HOST computer enable flag | 9 | 1 | c |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 5 |  |  | 0
33 | alrmmsgTA | Alarm message unit type | alarm message string with unit type information | 10 | 0 | c |  |  |  |  | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | 5 |  |  | 0
34 | alrmmsgnr | Alarm message number | alarm message string with unit number information | 10 | 1 | c |  |  |  |  | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | 5 |  |  | 0
35 | relstatus | Relay status | status of relays/potential free contacts ('0' = not activated, '1' = activated) | 10 | 2 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
36 | actualval | Actual counter value | actual value of counter |  | 0 | f |  | 0 | 0 | 3.4028e+38 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
37 | signinpsel | Signal input selection | signal input selection (' '=no value,'+'=pos value,'-'=neg value input) |  | 1 | c |  |  |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
38 | resinpsel | Reset input selection | external reset input enable/disable ('E'=enable,' '=disable) |  | 2 | c |  |  |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
39 | limit | <limit> | limit/batch for counter in sensor standard units |  | 3 | f |  |  | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 5 |  |  | 0
40 | delaytime | Delay time | delay time string in days,hours,minutes,seconds |  | 4 | c |  | 00000000 | 0 | 99235959 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
41 | duratntime | Duration time | duration time string in days,hours,minutes,seconds |  | 5 | c |  | 00000000 | 0 | 99235959 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
42 | vlvoutset | Valve output setting | valve output setting ('0'=do nothing, '1'=close valve) |  | 6 | c |  | 00000000 |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
43 | reloutset | Relay output setting | relay output setting ('L'=low, 'H'=high, 'P'=pulse (1 sec.)) |  | 7 | c |  | LLLLLLLL |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
44 | opermodeTA | Operation mode T/A | operation mode of T/A module |  | 8 | c |  | 0 | 0 | 9 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 8 | 0: OFF ¦ 1: A: MAX & RESP AUTO; T: UP TO LIMIT ¦ 2: A: MIN & RESP AUTO; T: UP AND REPEAT ¦ 3: A: MAX & RESP; T: DOWN FROM LIMIT ¦ 4: A: MIN & RESP; T: DOWN AND REPEAT ¦ 5: A: MAXIMUM ALARM; T: ALWAYS UP ¦ 6: A: MINIMUM ALARM ¦ 7: A: RESPONSE ALARM
45 | readunit | Readout unit | readout unit string |  | 9 | c |  | ln/min |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
46 | readfact | Readout factor | readout factor matching readout unit string |  | 10 | f |  | 1 | 1e-10 | 10000000000 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
47 | resetunit | Reset unit | reset unit command (1=reset T/A unit) |  | 12 | c |  |  | 0 | 1 | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 5 |  |  | 0
48 | TdValveDn | Valve differentiator down | valve output differentiation time constant downwards |  | 9 | f |  | 0.1 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 6 |  |  | 0
49 | TdValveUp | Valve differentiator up | valve output differentiation time constant upwards |  | 10 | f |  | 0.1 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 6 |  |  | 0
50 | TdSensorDn | Sensor differentiator down | sensor signal differentiation time constant downwards |  | 11 | f |  | 5.8 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 6 |  |  | 0
51 | TdSensorUp | Sensor differentiator up | sensor signal differentiation time constant upwards |  | 12 | f |  | 5.6 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 6 |  |  | 0
52 | CycleTime | Cycle time | cycle time * 10 msec. main loop signal processing | 114 | 12 | c | f | 6 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 6 |  |  | 0
53 | AnalogMode | Analog mode | analog mode selection for analog operation | 115 | 3 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 10 |  |  | 9 | 0: 0...5 Vdc operation ¦ 0: Analog input enabled ¦ 0: Analog output enabled ¦ 1: 0...10 Vdc operation ¦ 1: Analog input disabled ¦ 1: Analog output disabled ¦ 2: 0...20 mA operation ¦ 3: 4...20 mA operation ¦ 4: 15...20 mA operation
54 | VrefOutput | Reference voltage | reference voltage output signal for analog operation | 116 | 6 | i |  |  | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
55 | ValveOut | Valve output | valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc) | 114 | 1 | l |  | 0 | 0 | 16777215 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 8 |  |  | 0
56 | DynDispFct | Dynamic display factor | dynamic display factor for display filter (0=max, 1=min goes with par 57) | 117 | 1 | f |  | 0,001 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 6 |  |  | 0
57 | StaDispFct | Static display factor | static display factor for display filter (0=max, 1=min goes with par 56) | 117 | 2 | f |  | 0,000001 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 6 |  |  | 0
58 | CalMode | Calibration mode | calibration mode selection (not active until cntrlmode has been set to value 9) | 115 | 1 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 7 |  |  | 24 | 0: idle: no action ¦ 1: adc self calibration ¦ 2: dmfc ¦ 3: dmfc ¦ 4: dmfc ¦ 5: dmfc ¦ 6: dmfc ¦ 7: dmfc ¦ 8: dmfc ¦ 9: customer zero ¦ 10: adjust Vref output by connecting it to analog in ¦ 11: adjust analog out by connecting it to analog in ¦ 12: adjust valveoutput by connecting it to analog in ¦ 13: dmfc ¦ 14: dmfc ¦ 15: analog output = 0 % ¦ 16: analog output = 100 % ¦ 17: analog output = 50 % ¦ 18: factory zero ¦ 19: sensor differentiator (setpoint steps are needed!) ¦ 20: automatic sensor configuration ¦ 21: sensor temperature calibration ¦ 22: customer zero (no control mode 9 needed) ¦ 255: Error mode (result of previous cal mode)
59 | ValveOffst | Valve offset | valve offset: amount of DAC steps within 1 potmeter step | 116 | 7 | i |  | 61000 | -32767 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
60 | Monitor | Monitor mode | monitor: output signal (measure) selection for bus and analog output | 115 | 2 | c |  | 7 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 2 |  |  | 20 | 0: (filtered) setpoint ¦ 1: controller error input signal / raw sensor signal ¦ 2: controller output signal to valve ¦ 3: sensor signal slow ¦ 4: sensor signal slow filtered ¦ 5: linearization output ¦ 6: differentiator output ¦ 7: differentiator output filtered ¦ 8: normal sensor signal (Output) ¦ 9: analog input signal ¦ 10: power supply voltage ¦ 11: mass flow in display unit (normally ln/min) ¦ 12: volume flow in l/min ¦ 13: temperature in �C ¦ 14: pressure absolute in mbara ¦ 15: time in msec/frequency in Hz. ¦ 16: calibrated volume at actual sensor in ml ¦ 17: delta-P pressure in mbarg ¦ 18: atmospheric (barometer) pressure in mbara ¦ 19: mass flow in kg/min
61 | AlarmReg1 | Alarm register1 | alarm register containing warning flags | 114 | 2 | c |  |  |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 4 |  |  | 2 | 0: No diagnostics available in warning register ¦ 1: Diagnostics available in warning register
62 | AlarmReg2 | Alarm register2 | alarm register containing critical error flags | 114 | 3 | c |  |  |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 4 |  |  | 2 | 0: No diagnostics available in error register ¦ 1: Diagnostics available in error register
63 | CalRegZS1 | <CalRegZS1> | calibration register zero scale input 1 ADC | 116 | 1 | l |  | 210A7D | 0 | 16777215 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
64 | CalRegFS1 | <CalRegFS1> | calibration register full scale input 1 ADC | 116 | 2 | l |  | 52A513 | 0 | 16777215 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
65 | CalRegZS2 | <CalRegZS2> | calibration register zero scale input 2 ADC | 116 | 3 | l |  | 210A7D | 0 | 16777215 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
66 | CalRegFS2 | <CalRegFS2> | calibration register full scale input 2 ADC | 116 | 4 | l |  | 52A513 | 0 | 16777215 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
67 | ADCcntrReg | ADC control register | ADC control register | 114 | 4 | l |  | 18904E | 0 | 4294967295 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 12 | 0: ADC gain = 1x ¦ 0: ADC bipolar mode ¦ 0: Disable zero measure threshold ¦ 1: ADC gain = 2x ¦ 1: ADC unipolar mode ¦ 1: Enable zero measure threshold ¦ 2: ADC gain = 4x ¦ 3: ADC gain = 8x ¦ 4: ADC gain = 16x ¦ 5: ADC gain = 32x ¦ 6: ADC gain = 64x ¦ 7: ADC gain = 128x
68 | BridgePotm | Bridge potmeter | sensor bridge zero potmeter setting | 116 | 5 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
69 | AlarmEnble | <AlarmEnble> | broadcast alarm messsage enable flag | 115 | 4 | c |  | 1 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 4 |  |  | 2 | 0: disable ¦ 1: enable
70 | TestMode | Test mode | test mode selection (not active until cntrlmode has been set to value 5) | 115 | 5 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 4 |  |  | 13 | 0: idle; no action ¦ 1: uProcessor ¦ 2: IO ¦ 3: RAM ¦ 4: FRAM ¦ 5: ADC ¦ 6: DAC ¦ 7: sensor ¦ 8: valve drive circuit ¦ 9: Vref ¦ 10: FLOW-BUS ¦ 11: calibration ¦ 12: keyboard
71 | ChanSelect | <ADC channel select> | channel selection ADC | 115 | 6 | c |  | 1 | 1 | 32 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 2 | 1: AD channel 1 ¦ 2: AD channel 2
72 | ContrResp | Normal step controller response | controller response for normal steps (128=normal, <128=slower, >128=faster) | 114 | 5 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 |  |  | 0
73 | ErrInFilCo | Setpoint exponential smoothing filter | analog input filter constant (0=max, 1=min) | 117 | 3 | f |  | 1.0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 6 |  |  | 0
74 | ExpSmooCon | Sensor exponential smoothing filter | sensor input filter constant (0=max, 1=min) | 117 | 4 | f |  | 1.0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 6 |  |  | 0
75 | AnOutCorZS | Analog output zero scale | analog output correction factor zero scale (meas outp DSCM-A 0=0 other 32767=0) |  | 21 | i |  | 32767 | -32767 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
76 | AnOutCorFS | Analog output full scale | analog output correction factor full scale (meas outp 2000 = 1 * multiplication) |  | 22 | i |  | 2000 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
77 | AnInpCorZS | Analog input zero scale | analog input correction factor zero scale (ext setp DSCM-A 0=0 other 32767=0) |  | 23 | i |  | 32767 | -32767 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
78 | AnInpCorFS | Analog input full scale | analog input correction factor full scale (ext setp 2000 = 1 * multiplication) |  | 24 | i |  | 2000 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
79 | TuningMode | Tuning mode | (auto)tuning mode selection (not active until cntrlmode has been set to value 6) | 115 | 7 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 7 |  |  | 6 | 0: idle; no action ¦ 1: sensor ¦ 2: valve ¦ 3: Fuzzy controller normal operation ¦ 4: Fuzzy controller open at zero ¦ 5: PID controller
80 | DefVlvType | Valve default | valve type (needed for controlling behaviour) | 114 | 6 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 5 | 0: normally closed ¦ 1: normally opened ¦ 2: normally closed inverse controlled ¦ 3: normally opened inverse controlled ¦ 4: remain position
81 | GlobModify | Global modify | contains number(s) of changed processes for indirect polling (0xXX / 0xFF) | 0 | 19 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
82 | SpanCorr | Valve span correction factor | correction factor valve curve ratio high/low area | 114 | 7 | f |  | 0.1 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
83 | VlvCrvStps | Valve curve correction | Valve curve correction for controller (max. factor*0.1, flow where factor = 1) | 114 | 8 | c |  | 20,80 |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
84 | MemShipNor | <MemShipNor> | array with memberships for normal Fuzzy controller | 114 | 9 | c |  | 10,5000,10,5000 |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
85 | MemShipOpn | <MemShipOpn> | array with memberships for 0-open Fuzzy controller | 114 | 10 | c |  | 1,10000,3750,20000 |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
86 | IOStatus | IO status | IO status byte for jumper settings and LED signal modes | 114 | 11 | c |  | 4 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 12 | 20 |  | 8 | 1: read diagnostic jumper (no diagnostics, read/write) ¦ 1: not used ¦ 1: read micro switch (read/write) ¦ 1: diagnostic jumper set (read only) ¦ 1: initialization jumper set (read only) ¦ 1: analog jumper set (read only) ¦ 1: micro switch pressed (read only) ¦ 1: read analog jumper (use cntrlmode, read/write)
90 | DeviceType | Device type | (FLOW-BUS) device type information string | 113 | 1 | c |  | DMFC |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 13 |  |  | 0
91 | ModelNum | BHT model number | BHT model number information string | 113 | 2 | c |  | F201C-FA |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 13 |  |  | 0
92 | SerialNum | Serial number | serial number information string | 113 | 3 | c |  | SN999999A |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 13 |  |  | 0
93 | MfrConfig | Customer model number | customer model number information string | 113 | 4 | c |  | STANDARD |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 11 |  |  | 0
94 | BHT1 | BHT1 | special BHT parameter (to be changed by BHT only) | 118 | 1 | c |  | 01,01,95 |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 14 |  |  | 0
95 | BHT2 | BHT2 | special BHT parameter | 118 | 2 | i |  | 0 | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
96 | BHT3 | BHT3 | special BHT parameter | 118 | 3 | l |  | 0 | -3000000000 | 3000000000 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
97 | BHT4 | BHT4 | special BHT parameter | 118 | 4 | i |  | 0 | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
98 | BHT5 | BHT5 | special BHT parameter | 118 | 5 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
99 | BHT6 | BHT6 | special BHT parameter | 118 | 6 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
100 | BHT7 | BHT7 | special BHT parameter | 118 | 7 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
101 | BHT8 | BHT8 | special BHT parameter | 118 | 8 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
102 | BHT9 | BHT9 | special BHT parameter | 118 | 9 | l |  | 0 | -3000000000 | 3000000000 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
103 | BHT10 | BHT10 | special BHT parameter (to be changed by BHT only) | 118 | 10 | c |  | 0 | 0 | 1 | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 14 |  |  | 0
104 | PulseHight | Broadcast repeating time | broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve) | 114 | 16 | c |  | 128 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
105 | Version | Firmware version | revision number of firmware | 113 | 5 | c |  | VX.XX |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 13 | 20 |  | 0
106 | PressSensr | Pressure sensor type | type of pressure sensor | 115 | 9 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 20 |  |  | 10 | 0: delta-P 0..5" W.C. ¦ 1: delta-P 0...10" W.C. ¦ 2: absolute pressure 800-1200 mbar ¦ 3: absolute pressure 800-1100 mbar ¦ 4: delta-P -5...0 "W.C. ¦ 5: delta-P -10...0 "W.C. ¦ 6: delta-P -10...+10 "W.C. ¦ 7: delta-P 0...1 PSI ¦ 8: delta-P -1...0 PSI ¦ 9: absolute pressure 0-10 bar
107 | BaroPress | Barometer pressure | mbar atmospheric (central) barometer pressure | 116 | 8 | f |  | 1013.25 | 0 | 1200 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 20 |  |  | 0
108 | AnIn1CorZS | Sensor input zero scale | analog sensor signal input corr. factor zero scale (DSCM-A 0=0 other 32767=0) |  | 25 | i |  | 32767 | -32767 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
109 | AnIn1CorFS | Sensor input full scale | analog sensor signal input correction factor full scale (2000=1*multiplication) |  | 26 | i |  | 2000 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
110 | AnIn2CorZS | Reference voltage input zero scale | analog Vref input correction factor zero scale (DSCM-A 0=0 other 32767=0) |  | 27 | i |  | 32767 | -32767 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
111 | AnIn2CorFS | Reference voltage input full scale | analog Vref input correction factor full scale (2000=1*multiplication) |  | 28 | i |  | 2000 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
112 | AnOu1CorZS | Analog setpoint zero scale | analog setpoint output correction factor zero scale (DSCM-A 0=0 other 32767=0) |  | 29 | i |  | 32767 | -32767 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
113 | AnOu1CorFS | Analog setpoint full scale | analog setpoint output correction factor full scale (2000=1*multiplication) |  | 30 | i |  | 2000 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
114 | Reset | Reset | reset facilities (program/alarm/batchcounter) | 115 | 8 | c |  | 0 | 0 | 255 | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 12 |  |  | 6 | 0: no reset ¦ 1: reset counter value (no mode change) or common reset ¦ 2: reset alarm ¦ 3: restart batch counter ¦ 4: reset counter value (counter off) ¦ 5: Reset module (soft reset)
115 | UserTag | User tag | user definable alias string | 113 | 6 | c |  | USERTAG |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 11 |  |  | 0
116 | AlrmMaxLim | Alarm limit maximum | maximum limit for sensor signal to trigger alarm situation | 97 | 1 | i |  | 0 | 0 | 41600 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 15 |  |  | 0
117 | AlrmMinLim | Alarm limit minimum | minimum limit for sensor signal to trigger alarm situation | 97 | 2 | i |  | 0 | 0 | 41600 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 15 |  |  | 0
118 | AlrmMode | Alarm mode | alarm mode | 97 | 3 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 15 |  |  | 4 | 0: off ¦ 1: alarm on absolute limits ¦ 2: alarm on limits related to setpoint (response alarm) ¦ 3: alarm when instrument powers-up (eg. after power-down)
119 | AlrmOutMod | Alarm output mode | alarm relais activity mode during alarm situation | 97 | 4 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 15 |  |  | 3 | 0: no relais activity at alarm ¦ 1: relais pulses until reset ¦ 2: relais activated until reset
120 | AlrmStpMod | Alarm setpoint mode | setpoint change enable during alarm situation | 97 | 5 | c |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 15 |  |  | 2 | 0: no setpoint change at alarm ¦ 1: new/safe setpoint at alarm enabled (set at par 121)
121 | AlrmNwSetp | Alarm new setpoint | new/safe setpoint during alarm situation (until reset) | 97 | 6 | i |  | 0 | 0 | 32767 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 15 |  |  | 0
122 | CntrValue | Counter value | actual counter value | 104 | 1 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
123 | CntrUnit | Counter unit index | pointer to counter unit (classic counter unit table) | 104 | 2 | c |  | 0 | 0 | 31 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
124 | CntrLimit | Counter limit | counter limit/batch | 104 | 3 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
125 | CntrOutMod | Counter output mode | counter relais activity mode when limit/batch has been reached | 104 | 4 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 3 | 0: no relais activity at batch limit ¦ 1: relais pulses after reaching batch limit until reset ¦ 2: relais activated after reaching batch limit until reset
126 | CntrStpMod | Counter setpoint mode | setpoint change enable during counter limit/batch situation (until reset) | 104 | 5 | c |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 2 | 0: setpoint change at batch limit disabled ¦ 1: setpoint change at batch limit enabled
127 | CntrNwSetp | Counter new setpoint | new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%) | 104 | 6 | i |  | 0 | 0 | 32767 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
128 | CntrUntstr | Counter unit | counter readout unit (informative only for older devices) | 104 | 7 | c |  | ln |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
129 | capunitstr | Capacity unit | capacity readout unit (informative only for older devices) |  | 31 | c |  | ln/min |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 | 19 |  | 0
130 | CntrMode | Counter mode | counter mode | 104 | 8 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 3 | 0: off ¦ 1: counting upwards continuously ¦ 2: counting up to limit (batchcounter)
131 | HwRev | Minimum hardware revision | minimum required hardware revision level for firmware version | 113 | 7 | c |  | VX.XX |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 13 |  |  | 0
132 | RCreadfact | <RCreadfact> | readout factor for direct reading (changes with readunit: local on module, R.O.) |  | 1 | f |  | 1.0 | 1e-10 | 10000000000 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
133 | channumber | <channumber> | channel number for operation |  | 2 | c |  | 1 | 1 | 120 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
134 | masterchan | <masterchan> | master channel for master-slave operation |  | 3 | c |  | 0 | 0 | 120 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
135 | RCslavefct | <RCslavefct> | RC slave factor |  | 4 | i |  | 32000 | 0 | 32000 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
136 | inputnode | <inputnode> | physical node address for channel number |  | 5 | c |  | 3 | 0 | 128 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 17 |  |  | 0
137 | inputproc | <inputproc> | physical process for channel number |  | 6 | c |  | 1 | 0 | 128 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 17 |  |  | 0
138 | RCreadunit | <RCreadunit> | readout unit for direct reading (local variable on module: read only) |  | 7 | c |  | ln/min |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 17 |  |  | 0
139 | SlaveFact% | Slave factor | slave factor for master slave control (setp = master output * slave factor) | 33 | 1 | f |  | 100.0 | 0 | 500 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 18 |  |  | 0
140 | VrefInput | Reference voltage input | reference voltage input for setpoint signal | 33 | 2 | i |  | 0 | 0 | 65535 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | 18 |  |  | 0
141 | RespStable | Stable situation controller response | controller response when controller is stable: ¦measure-setpoint¦ < 2% | 114 | 17 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 |  |  | 0
142 | temperatur | Temperature | absolute temperature in degrees Celsius | 33 | 7 | f |  | 20 | -250 | 500 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 19 | 13 |  | 0
143 | pressure | Pressure | absolute pressure in mbar (FPP, FRM, FTM) or bar | 33 | 8 | f |  | 1013.25 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 19 | 13 |  | 0
144 | time | Time | time in milliseconds | 33 | 9 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
145 | calvolume | Calibrated volume | calibrated volume in litres | 33 | 10 | f |  | 50 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 19 |  |  | 0
146 | sensornr | Sensor number | pointer to sensor number in calibration tube (FPP) |  | 16 | c |  | 0 | 0 | 4 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
147 | rangeselct | Range select | Piston Prover operation mode (write) and status information (read back) | 115 | 10 | c |  | 0 | 0 | 99 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 20 | 19 |  | 13 | 0: calibration ready/stop ¦ 1: run calibration until stopsensor 1/select range 1 ¦ 2: run calibration until stopsensor 2/select range 2 ¦ 3: run calibration until stopsensor 3/select range 3 ¦ 4: run calibration until stopsensor 4/select range 4 ¦ 5: run calibration and select range 5 ¦ 9: run calibration with automatic range selection ¦ 19: run until stopsensor 1 until 3 values between limit ¦ 29: run until stopsensor 2 until 3 values between limit ¦ 39: run until stopsensor 3 until 3 values between limit ¦ 49: run until stopsensor 4 until 3 values between limit ¦ 59: run and select range 5 until 3 values between limit ¦ 99: run with auto-select + 3 values between limit
148 | TimeOut | Time out | maximum admitted duration time for specific procedure (in 100 ms) |  | 2 | i |  | 0 | 0 | 30000 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 20 |  |  | 0
149 | frequency | Frequency | frequency in Hz | 33 | 9 | f |  | 0 | 0 | 100000 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 21 |  |  | 0
150 | imp/m3 | Impulses/m3 | For FRM and FTM imp/m3 and for FCM imp/kg | 33 | 10 | f |  | 42773.4 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 20 |  |  | 0
151 | RefVolFlow | Normal volume flow | volume flow referenced to normal conditions i.e. 0 �C, 1013.25 hPa(a) in ln/min | 33 | 5 | f |  | 0 | 0 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
152 | volumeflow | Volume flow | volume flow at actual conditions in l/min | 33 | 6 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
153 | delta-p | Delta-p | relative pressure between atmosphere and sensor position | 33 | 11 | f |  | 0 | -100000 | 100000 | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
154 | scalefact | <scalefact> | scaling factor (multiplication) for readout on display (for optimal resolution) | 33 | 13 | i |  | 1 | 1 | 10000 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 21 |  |  | 0
155 | sensorname | Sensor name | label with information about stop sensor |  | 17 | c |  | SENSOR0 |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 19 |  |  | 0
156 | RstAlarmEn | Reset alarm enable | enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic | 97 | 9 | c |  | 15 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 15 |  |  | 16 | 0: no reset possible ¦ 1: reset: keyboard ¦ 2: reset: external ¦ 3: reset: keyboard or external ¦ 4: reset: FLOW-BUS ¦ 5: reset: FLOW-BUS or keyboard ¦ 6: reset: FLOW-BUS or external ¦ 7: reset: FLOW-BUS or keyboard or external ¦ 8: reset: automatic ¦ 9: reset: automatic or keyboard ¦ 10: reset: automatic or external ¦ 11: reset: automatic or keyboard or external ¦ 12: reset: automatic or FLOW-BUS ¦ 13: reset: automatic or FLOW-BUS or keyboard ¦ 14: reset: automatic or FLOW-BUS or external ¦ 15: reset: automatic or FLOW-BUS or keyboard or external
157 | RstCountEn | Reset counter enable | enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic | 104 | 9 | c |  | 7 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 16 |  |  | 16 | 0: no reset possible ¦ 1: reset: keyboard ¦ 2: reset: external ¦ 3: reset: keyboard or external ¦ 4: reset: FLOW-BUS ¦ 5: reset: FLOW-BUS or keyboard ¦ 6: reset: FLOW-BUS or external ¦ 7: reset: FLOW-BUS or keyboard or external ¦ 8: reset: automatic ¦ 9: reset: automatic or keyboard ¦ 10: reset: automatic or external ¦ 11: reset: automatic or keyboard or external ¦ 12: reset: automatic or FLOW-BUS ¦ 13: reset: automatic or FLOW-BUS or keyboard ¦ 14: reset: automatic or FLOW-BUS or external ¦ 15: reset: automatic or FLOW-BUS or keyboard or external
158 | MasterNode | Master node | node number of master instrument output signal for a slave | 33 | 14 | c |  | 3 | 1 | 128 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 18 |  |  | 0
159 | MasterProc | Master process | process number of master instrument output signal for a slave | 33 | 15 | c |  | 1 | 1 | 128 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 18 |  |  | 0
160 | InstrNode | Remote instrument node | node number of instrument to be operated by another module (keyboard/display) | 33 | 16 | c |  | 3 | 1 | 128 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 18 |  |  | 0
161 | InstrProc | Remote instrument process | process number of instrument to be operated by another module (keyboard/display) | 33 | 17 | c |  | 1 | 1 | 128 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 18 |  |  | 0
162 | RangeMin | Minimum custom range | Mnimum value at 0% for special user readout unit | 33 | 18 | f |  | 0.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
163 | RangeMax | Maximum custom range | Maximum value at 100% for special user readout unit | 33 | 20 | f |  | 100.0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
164 | Relay/TTL | Relay/TTL output | Relay/TTL output setting (disabled when used by alarm or counter) | 115 | 11 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 12 |  |  | 0
165 | RespOpen0% | Open from zero controller response | Controller response when valve opens from zero | 114 | 18 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 |  |  | 0
166 | ContrType | Controller features | Controller settings for special purpose | 114 | 20 | c |  | 1 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 16 | 0: valve in normal position after startup ¦ 0: open from zero with PID output to valve ¦ 0: fixed monitor output signal ¦ 0: voltage drift compensation for valve output turned on ¦ 0: auto slope disabled ¦ 0: automatic correction for valve open turned on ¦ 0: controller special mode (valve output steps) turned off ¦ 0: valve overshoot protection turned off ¦ 1: valve in safe position after startup ¦ 1: open from zero with ramp output to valve ¦ 1: monitor output changed at setpoint steps ¦ 1: voltage drift compensation for valve output turned off ¦ 1: auto slope enabled for pilot valves ¦ 1: automatic correction for valve open turned off ¦ 1: controller special mode (valve output steps) turned on ¦ 1: valve overshoot protection turned on
167 | PIDKp | PID-Kp | PID factor Kp | 114 | 21 | f |  | 10 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 | 3 |  | 0
168 | PIDTi | PID-Ti | PID factor Ti | 114 | 22 | f |  | 0.05 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 |  |  | 0
169 | PIDTd | PID-Td | PID factor Td | 114 | 23 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 |  |  | 0
170 | Density | Density | Density of selected fluid in kg/m3 | 33 | 21 | f |  | 1.293 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
171 | CalCertNr | Calibration certificate | Number of calibration certificate (last basic calibration) | 113 | 8 | c |  |  |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 13 | 3 |  | 0
172 | CalDate | Calibration date | Date of last (basic) calibration | 113 | 9 | c |  | 19991231 |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 13 | 3 |  | 0
173 | ServiceNr | Service number | Service number for repair/rebuilding/recalibration | 113 | 10 | c |  | 000000 |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 13 |  |  | 0
174 | ServDate | Service date | Date of last service action | 113 | 11 | c |  | 19991231 |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 13 |  |  | 0
175 | IdentNr | Identification number | Identification number (type) of instrument/device | 113 | 12 | c |  | 7 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 13 |  |  | 43 | 0: UFO?: Unidentified FLOW-BUS Object ¦ 1: RS232/FLOW-BUS interface ¦ 2: PC(ISA) interface ¦ 3: ADDA4 (4 channels) ¦ 4: R/C-module, 32 channels ¦ 5: T/A-module ¦ 6: ADDA1: 1 channel ADDA converter module ¦ 7: DMFC: digital mass flow controller ¦ 8: DMFM: digital mass flow meter ¦ 9: DEPC: digital electronic pressure controller ¦ 10: DEPM: digital electronic pressure meter ¦ 11: ACT: single actuator ¦ 12: DLFC: digital liquid flow controller ¦ 13: DLFM: digital liquid flow meter ¦ 14: DSCM-A: digital single channel module for analog instruments ¦ 15: DSCM-D: digital single channel module for digital instr. ¦ 16: FRM: FLOW-BUS rotor meter (calibration-instrument) ¦ 17: FTM: FLOW-BUS turbine meter (calibration-instrument) ¦ 18: FPP: FLOW-BUS piston prover/tube (calibration-instrument) ¦ 19: F/A-module: special version of T/A-module ¦ 20: DSCM-E: evaporator controller module (single channel) ¦ 21: DSCM-C: digital single channel module for calibrators ¦ 22: DDCM-A: digital dual channel module for analog instruments ¦ 23: DMCM-D: digital multi channel module for digital instruments ¦ 24: Profibus-DP/FLOW-BUS interface module ¦ 25: FLOW-BUS Coriolis Meter ¦ 26: FBI: FLOW-BUS Balance Interface ¦ 27: CORIFC: CoriFlow Controller ¦ 28: CORIFM: CoriFlow Meter ¦ 29: FICC: FLOW-BUS Interface Climate Control ¦ 30: IFI: Instrument FLOW-BUS Interface ¦ 31: KFI: Keithley FLOW-BUS Interface ¦ 32: FSI: FLOW-BUS Switch Interface ¦ 33: MSCI: Multi-Sensor/Confroller Interface ¦ 34: APP-D: Active Piston Prover ¦ 35: LFI: Leaktester FLOW-BUS Interface ¦ 36: Digital Batch Flow Controller ¦ 37: Digital PID Controller ¦ 38: Bus Gateway ¦ 39: DTC: digital temperature controller ¦ 40: DTM: digital temperature meter ¦ 41: DC: digital controller ¦ 42: DM: digital meter
176 | BHT11 | BHT11 | special BHT parameter (to be changed by BHT only) | 118 | 11 | c |  | 0 | 0 | 255 | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 14 |  |  | 0
177 | PowerMode | Power mode | power supply indication in Vdc | 115 | 12 | c |  | 0 | 0 | 50 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 12 |  |  | 0
178 | Pupstream | Pressure inlet | pressure inlet (upstream) of fluid in bara (for first fluidnr only) | 113 | 13 | f |  | 3 | -100000 | 100000 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 | 3 |  | 0
179 | Pdownstrm | Pressure outlet | pressure outlet (downstream) of fluid in bara (for first fluidnr only) | 113 | 14 | f |  | 1 | -100000 | 100000 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 | 3 |  | 0
180 | Orifice | Orifice | orifice diameter in mm | 113 | 15 | f |  | 1 | 0 | 1000 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 |  |  | 0
181 | FluidTemp | Fluid temperature | temperature of fluid through instrument (for first fluidnr only) | 113 | 16 | f |  | 20 | -273.15 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 | 3 |  | 0
182 | AlrmDelay | Alarm delay | time alarm and reset action will be delayed when alarm limit has been exceeded | 97 | 7 | c |  | 3 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 15 |  |  | 0
183 | capacity0% | Capacity 0% | capacity of instrument at zero 0% in sensor base units (mostly equal to zero) | 33 | 22 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
184 | NumOfChan | Number of channels | number of instrument channels available for this device | 0 | 18 | c |  | 1 | 0 | 120 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
185 | DeviceFunc | Device function | function of device | 0 | 20 | c |  | 5 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 9 | 0: Unknown ¦ 1: Interface ¦ 2: ADDA ¦ 3: Operator ¦ 4: Supervisor (totalizer/alarm) ¦ 5: Controller ¦ 6: Meter ¦ 7: Special ¦ 8: (Protocol) converter
186 | ScanChan | Scan channel | Channel number to scan with real time information (to be set once) | 123 | 1 | c |  | 1 | 1 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 0
187 | ScanPar | Scan parameter | Parameter number to scan with real time information (to be set once) | 123 | 3 | c |  | 8 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 0
188 | ScanTime | Scan time | Scan interval time in msec between two samples (to be set once) | 123 | 4 | i |  | 50 | 0 | 65535 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 0
189 | ScanData | Scan data | Scanned data with time label (can be readout event by event) | 123 | 10 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 22 | 4 |  | 0
190 | ValveOpen | Valve open | First-step offset current/voltage for valve when opening from 0% | 114 | 24 | f |  | 0.04 | 0 | 24 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 |  |  | 0
191 | NrOfRuns | Number of runs | Amount of runs of a piston prover (0 = stability check) | 115 | 13 | c |  | 1 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 20 |  |  | 0
192 | MinProTime | Minimum process time | Minimum process time of a piston prover in 0.1 seconds | 115 | 14 | c |  | 10 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 20 |  |  | 0
193 | LeakRate | Leak rate | Leak rate piston prover/piston tare value multiplier | 116 | 9 | f |  | 0.0001 | 0 | 3 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 20 |  |  | 0
194 | ModeInfReq | Mode info request | Sets instr. in info mode for 1 read-cycle to check available parameter options | 115 | 15 | c |  |  |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 12 |  |  | 0
195 | ModeInfOpt | Mode info option list | Gives info about possible values of a mode in an array as result of ModeInfo req | 115 | 16 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
196 | ModeInfDes | Mode info option description | Gives description about one of the mode options | 115 | 17 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
197 | CalType | Calibrations options | Enables/disables options for calibration device (8 bits for 8 options) | 115 | 18 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 20 |  |  | 4 | 0: Automatic capacity setting for optimal resolution ¦ 0: Barometer value input via parameter 107: BaroPress ¦ 1: Manual capacity setting for optimal resolution ¦ 1: Barometer is master; input automatically from master
198 | MassFlow | Mass flow | Real mass flow in kg/min | 33 | 4 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 20 |  |  | 0
199 | BusAddress | Fieldbus1 address | Fieldbus1 address (main interface) | 125 | 10 | c |  | 2 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 23 |  |  | 0
200 | InterfConf | Interface configuration | Configuration setting for interface to other bus-systems | 125 | 3 | c |  | 1 | 0 | 3 | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | 23 |  |  | 4 | 0: Configuration A: 14 ch. Standard parms. with network scan ¦ 1: Configuration B: 14 ch. Standard parms with fixed chan list ¦ 2: Configuration C: 7 ch. Extended parms with fixed chan list ¦ 3: Configuration D: 11 ch. Extended parms with network scan
201 | Baudrate | Fieldbus1 baudrate | Fieldbus1 baudrate (main interface) | 125 | 9 | l |  | 12000000 | 0 | 10000000000 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 23 |  |  | 0
202 | BusDiagnos | Fieldbus1 diagnostic string | Fieldbus1 diagnose string (main interface) | 125 | 20 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 23 |  |  | 0
203 | NrOfVanes | Number of vanes | Number of vanes for use in a rotor meter | 115 | 22 | c |  | 10 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 20 |  |  | 0
204 | Fieldbus | Fieldbus1 name | Fieldbus1 name (main interface) | 125 | 21 | c |  | FLOW-BUS |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 23 |  |  | 0
205 | fMeasure | fMeasure | measured value for direct reading (in capunits, max.= capacity) | 33 | 0 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 2 |  |  | 0
206 | fSetpoint | fSetpoint | setpoint: wanted value for direct reading (in capunits, max.= capacity) | 33 | 3 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 2 | 18 |  | 0
207 | Mass | Mass | Mass in g | 33 | 23 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 20 |  |  | 0
208 | Mstatus | Manufacturer status register | Manufacturer Status register (64 diagnostic bits) | 119 | 1 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 2 | 0: No diagnostics available in manufacturer status register ¦ 1: Diagnostics available in manufacturer status register
209 | Mwarning | Manufacturer warning register | Manufacturer Warning register (64 diagnostic bits) | 119 | 2 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 2 | 0: No diagnostics available in manufacturer warning register ¦ 1: Diagnostics available in manufacturer warning register
210 | Merror | Manufacturer error register | Manufacturer Error register (64 diagnostic bits) | 119 | 3 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 2 | 0: No diagnostics available in manufacturer error register ¦ 1: Diagnostics available in manufacturer error register
211 | DiagHist | Diagnostic history string | Diagnostic history string (contains history of diag codes) | 119 | 4 | c |  |  |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 0
212 | DiagMode | Diagnostic mode | Diagnostic mode (0 = diagnostics off, 1 = diagnostics on) | 119 | 5 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 4 |  |  | 2 | 0: Debug mode off ¦ 1: Debug mode on
213 | MStatEnabl | Manufacturer status enable | Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all) | 119 | 6 | c |  | 0 | 0 | 255 | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 |  |  | 4 | 0: set status bit (range 0�127) ¦ 127: set status bit (range 0�127) ¦ 254: clear all status bits ¦ 255: set all status bits
214 | AnOutZA | Analog output zero adjust | Analog measure output, zero adjust | 116 | 21 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
215 | AnOutSA | Analog output span adjust | Analog measure output, span adjust | 116 | 22 | f |  | 1 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
216 | AnInZA | Analog input zero adjust | Analog setpoint input, zero adjust | 116 | 23 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
217 | AnInSA | Analog input span adjust | Analog setpoint input, span adjust | 116 | 24 | f |  | 1 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
218 | SensInZA | Sensor input zero adjust | Sensor input, zero adjust | 116 | 25 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
219 | SensInSA | Sensor input span adjust | Sensor input, span adjust | 116 | 26 | f |  | 1 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
220 | TempInZA | Temperature input zero adjust | Sensor temperature input, zero adjust | 116 | 27 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
221 | TempInSA | Temperature input span adjust | Sensor temperature input, span adjust | 116 | 28 | f |  | 1 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
222 | ExpSmooAd | Adaptive smoothing factor | Sensor input filter adapt setting | 117 | 5 | f |  | 1.0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 6 |  |  | 0
223 | SlopeSetp | Slope setpoint step | Slope setpoint step. Setpoint step for the given slopetime | 33 | 24 | i |  | 32000 | 0 | 32000 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 18 |  |  | 0
224 | FilterLen | Filter length | Number of samples for average filter | 117 | 6 | i |  | 1 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 6 |  |  | 0
225 | fAccuracy | Absolute accuracy | Actual accuracy in current unit | 33 | 25 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 2 |  |  | 0
226 | LookI | Lookup table index | Lookup table for linearisation index (x and y direction) | 33 | 26 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
227 | LookX | Lookup table X | Lookup table for linearisation x | 33 | 27 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
228 | LookY | Lookup table Y | Lookup table for linearisation y | 33 | 28 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
229 | LookTempI | Lookup table temperature index | Lookup table for linearisation at certain temperature index (z direction) | 33 | 29 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 3 |  |  | 0
230 | LookTemp | Lookup table temperature | Lookup table for linearisation at certain temperature (z) | 33 | 30 | f |  | 0 | -273.15 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
231 | ValveMax | Valve maximum | Maximum current/voltage for valve | 114 | 25 | f |  | 0.2 | 0 | 24 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
232 | ValveMode | Valve mode | Valve output mode selection | 114 | 26 | c |  | 1 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 3 | 0: voltage drive mode ¦ 1: current drive mode ¦ 2: 0V (for PIO use)
233 | VlvOpenCor | Valve open correction | Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96) | 114 | 27 | f |  | 0.96 | 0 | 1.5 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
234 | VlvZeroHld | Valve zero hold | Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8) | 114 | 28 | f |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
235 | ValveSlope | Valve slope | Valve slope time (seconds) | 114 | 29 | f |  | 0.009 | 0 | 50 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
236 | IFIData | IFI data | IFI data dump protocol communication string | 0 | 21 | c |  | 0 |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 23 |  |  | 0
237 | RangeUsed | Range used | Piston Prover information about used sensors | 115 | 20 | c |  | 0 | 0 | 99 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 20 | 19 |  | 0
238 | FldSetProp | Fluidset properties | Fluidset properties | 33 | 31 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 3 |  |  | 6 | 0: Fluidset is disabled ¦ 0: Fluidset is not factory set ¦ 0: Fluidset is not calibrated on actual gas ¦ 1: Fluidset is enabled ¦ 2: Fluidset is factory set ¦ 4: Fluidset is calibrated on actual gas
239 | LUnitType | Lookup table unit type index | Lookup table unit type | 33 | 12 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 3 |  |  | 0
240 | LUnTypName | Lookup table unit type | Lookup table unit type name | 33 | 13 | c |  |  |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 3 |  |  | 0
241 | LUnit | Lookup table unit index | Lookup table unit (unit LUTy) | 33 | 16 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 3 |  |  | 0
242 | LUnitName | Lookup table unit | Lookup table unit name | 33 | 17 | c |  | kg/s |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | 3 |  |  | 0
243 | CUnitType | Capacity unit type index | Capacity (readout) unit type |  | 29 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
244 | CUnTypName | Capacity unit type | Capacity (readout) unit type name |  | 30 | c |  |  |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
245 | CUnTypTemp | Capacity unit type temperature | Capacity (readout) unit type temperature (�C) | 33 | 10 | f |  | 0 | -273.15 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
246 | CUnTypPres | Capacity unit type pressure | Capacity (readout) unit type pressure (bar (a)) | 33 | 11 | f |  | 1 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
247 | CapMin | Capacity minimum | Minimum capacity in output capacity units |  | 27 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
248 | CapMax | Capacity maximum | Maximum capacity in output capacity units |  | 28 | f |  | 1 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
249 | FormulaTyp | Formula type | Formula type needed for conversion | 113 | 17 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
250 | HeatCap | Heat capacity | Heat capacity (Cp) (sensor conditions) | 113 | 18 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
251 | ThermCond | Thermal conductivity | Thermal conductivity (sensor conditions) | 113 | 20 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
252 | Viscosity | Viscosity | Dynamic viscosity (fluid conditions) | 113 | 21 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
253 | NormMasFlw | Standard flow | Standard mass flow in ln/min (20�C, 1.01325 bar (a)) air or g/h H2O equivalent | 113 | 22 | f |  | 1 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
254 | Kspeed | Controller speed | Controller speed factor (gain) | 114 | 30 | f |  | 1 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 8 | 3 |  | 0
255 | SensorCode | Sensor code | Sensor code | 113 | 23 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
256 | SensorRevC | Sensor configuration code | Sensor configuration code | 113 | 24 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
257 | RestrCode | Restriction code | Restriction code | 113 | 25 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
258 | RestrRevC | Restriction configurator code | Restriction configuration code | 113 | 26 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
259 | RestrNxP | Restriction NxP | Restriction NxP (proportional to air equivalent capacity of LFE) | 113 | 27 | l |  | 0 | 0 | 2147483648 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
260 | Seals | Seals information | Seals information (1st byte = other, 2nd = plunger seal) | 113 | 28 | c |  | V,V | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
261 | ValveCode | Valve code | Valve code | 113 | 29 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
262 | ValveRevC | Valve configuration code | Valve configuration code | 113 | 30 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
263 | InstrProp | Instrument properties | Instrument properties | 113 | 31 | l |  | 0 | 0 | 2147483648 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
264 | LookFreqI | Lookup table frequency index | Lookup table for frequency index | 116 | 10 | c |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
265 | LFFreq | Lookup table frequency frequency | Lookup table for frequency frequency | 116 | 11 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
266 | LFTemp | Lookup table frequency temperature | Lookup table for frequency temperature | 116 | 12 | f |  | 0 | -273.15 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
267 | LFDensity | Lookup table frequency density | Lookup table for frequency density | 116 | 13 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
268 | LFSpanAdj | Lookup table frequency span adjust | Lookup table for frequency span adjust | 116 | 14 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
269 | CUnit | Capacity unit index (ext) | Capacity (readout) unit index (extended unit table) | 65 | 15 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 | 19 |  | 0
270 | DensityAct | Density actual | Actual density, measured by instrument | 116 | 15 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
271 | RestrMeas | Measured restriction | Measured restriction | 116 | 18 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 |  |  | 0
272 | TempPotm | Temperature potmeter | Potmeter for sensor temperature compensation | 116 | 8 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
273 | TempGain | Temperature potmeter gain | Gain for sensor temperature compensation | 116 | 9 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
274 | CntrCConv | Counter controller overrun correction | Counter controller overrun correction | 104 | 10 | f |  | 1 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
275 | CntrCGain | Counter controller gain | Counter controller gain | 104 | 11 | f |  | 1 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
276 | FluidSub | Sub fluid number | Sub fluid number | 65 | 1 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
277 | FreqTCor | Temperature compensation factor | Temperature compensation factor | 116 | 17 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
278 | DSPRegI | DSP register address | DSP register address | 116 | 29 | l | c | 0 | 0 | 4294967295 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 13 |  |  | 0
279 | DSPRegLng | DSP register long | DSP register long | 116 | 30 | l |  | 0 | -4294967295 | 4294967295 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 |  |  | 0
280 | DSPRegFlt | DSP register floating point | DSP register floating point | 116 | 30 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 |  |  | 0
281 | DSPRegInt | DSP register integer | DSP register integer | 116 | 31 | i |  | 0 | -65535 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 13 |  |  | 0
282 | StdDev | Standard deviation | Standard deviation | 121 | 0 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
283 | MeasStatus | Measurement status | Measurement status | 121 | 1 | i |  | 0 | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
284 | MStopCrit | Measurement stop criteria | Measurement stop criteria | 121 | 2 | i |  | 0 | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 19 |  |  | 0
285 | MTimeOut | Measurement time out | Measurement time out | 121 | 3 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 20 |  |  | 0
286 | MMaxNrRuns | Maximum number of runs | Measurement maximum number of runs | 121 | 4 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 20 |  |  | 0
287 | MMinStdDev | Minimum standard deviation | Measurement minimum standard deviation | 121 | 5 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 20 |  |  | 0
288 | IOSwitchSt | IO switch status | IO status for switches | 114 | 31 | l |  | 0 | 0 | 4294967295 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 10 |  |  | 0
295 | SensBridge | Sensor bridge settings | Sensor bridge settings | 65 | 21 | i |  | 513 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 9 |  |  | 5 | 1: Bridge on ¦ 1: 3 windings C ¦ 1: 3 windings D ¦ 1: Sensor protection enabled ¦ 1: Automatic sensor configuration on
296 | SensCurren | Sensor bridge current | Sensor bridge current (0..0.02A) | 65 | 22 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 9 |  |  | 0
297 | SensResist | Sensor resistance | Sensor resistance (Ohm) | 65 | 23 | f |  | 0 | 0 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 9 |  |  | 0
298 | SensVolt | Sensor bridge voltage | Sensor bridge voltage (0..11.5V) | 65 | 24 | f |  | 0 | 0 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 9 |  |  | 0
299 | SensName | Sensor group name | Sensor name (based on sensor detection / sensor resistance) | 65 | 25 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 9 |  |  | 0
300 | SensCalTmp | Sensor calibration temperature | Sensor temperature at calibration | 116 | 20 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
301 | ValveSafe | Valve safe state | Valve safe state | 115 | 31 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 8 |  |  | 6 | 0: 0 mA ¦ 1: max mA ¦ 2: Close ¦ 3: Open ¦ 4: Idle ¦ 5: Value (for DeviceNet only)
302 | CnUnitType | Counter unit type index | Counter unit type | 104 | 12 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
303 | CnUnTypNam | Counter unit type | Counter unit type name | 104 | 13 | c |  |  |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
304 | CnUnit | Counter unit index (ext) | Counter unit index (extended counter unit table) | 104 | 14 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
305 | Bus1Select | Fieldbus1 selection | Fieldbus1 select (main interface) | 125 | 8 | c |  | 255 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 8 | 0: FLOW-BUS ¦ 1: Modbus ¦ 2: ProPar ¦ 10: DeviceNet ¦ 11: EtherCAT ¦ 13: PROFIBUS DP ¦ 14: PROFINET ¦ 255: NoBus
306 | Bus1Medium | Fieldbus1 medium | Fieldbus1 medium id (main interface) | 125 | 11 | c |  | 1 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 2 | 0: RS232 ¦ 1: RS485
307 | Bus2Mode | Fieldbus2 mode | Fieldbus2 mode | 124 | 7 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 2 | 0: Normal ¦ 1: Config mode
308 | Bus2Select | Fieldbus2 selection | Fieldbus2 select | 124 | 8 | c |  | 2 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 3 | 0: FLOW-BUS ¦ 1: Modbus ¦ 2: ProPar
309 | Bus2Addr | Fieldbus2 address | Fieldbus2 address | 124 | 10 | c |  | 2 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 23 |  |  | 0
310 | Bus2Baud | Fieldbus2 baudrate | Fieldbus2 baudrate | 124 | 9 | l |  | 12000000 | 0 | 10000000000 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 23 |  |  | 0
311 | Bus2Medium | Fieldbus2 medium | Fieldbus2 medium id | 124 | 11 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 2 | 0: RS232 ¦ 1: RS485
312 | Bus2Diagn | Fieldbus2 diagnostics | Fieldbus2 diagnose string | 124 | 20 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 23 |  |  | 0
313 | Bus2Name | Fieldbus2 name | Fieldbus2 name | 124 | 21 | c |  | Propar |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 23 |  |  | 0
314 | PIOChSel | PIO channel selection | PIO channel selection | 120 | 0 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 3 | 0: Analog input ¦ 1: Analog output ¦ 2: General purpose in-/output
315 | PIOParm | PIO parameter | PIO parameter connected to i/o channel (=process number*256+parameter number) | 120 | 2 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
316 | PIOFilter | PIO input/output filter | PIO input/output filter constant (0..1, 0=max, 1=min) | 120 | 6 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
317 | PIOPrmCap0 | PIO parameter capacity 0% | PIO parameter capacity, 0% value | 120 | 7 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
318 | PIOPrmCap | PIO parameter capacity 100% | PIO parameter capacity, 100% value | 120 | 3 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
319 | PIOCfgSel | PIO configuration selection | PIO i/o channel configuration selection | 120 | 1 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 10 | 0: Voltage output (0..10 V) ¦ 1: Current output (0..20 mA) ¦ 2: Digital output ¦ 3: Frequency output ¦ 4: Duty cycle output (20 kHz) ¦ 5: Digital pulse output ¦ 6: Voltage input (0..10 V) ¦ 7: Current input (0..20 mA) ¦ 8: Digital input ¦ 255: Disabled
320 | PIOAnZA | PIO analog zero adjust | PIO analog input/output, zero adjust value | 120 | 4 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
321 | PIOAnSA | PIO analog span adjust | PIO analog input/output, span adjust factor | 120 | 5 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
322 | PIOHwCapMx | PIO hardware capacity max | PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.) | 120 | 8 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
323 | PIOCapSel | PIO capacity set selection | PIO capacity set selection | 120 | 9 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
324 | PIOHwCap0 | PIO hardware capacity 0% | PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.) | 120 | 10 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
325 | PIOHwCap | PIO hardware capacity 100% | PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.) | 120 | 11 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 10 |  |  | 0
326 | HardwId | Hardware platform id | Hardware platform identification number | 0 | 6 | i |  | 0 | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 12 |  |  | 0
327 | HardwSubId | Hardware platform sub id | Hardware platform sub identification number | 0 | 7 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 12 |  |  | 0
328 | TempBaud | Temporary baudrate | Temporary volatile fieldbus2 baudrate | 124 | 31 | l |  | 0 | 0 | 10000000000 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 23 |  |  | 0
329 | SetMonitor | Setpoint monitor mode | Monitor mode for setpoint | 115 | 23 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 2 |  |  | 3 | 0: Setpoint ¦ 1: Filtered setpoint ¦ 2: Setpoint after linear slope
330 | BHT12 | BHT12 | Special BHT parameter | 118 | 12 | l |  | 0 | 0 | 10000000000 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 14 |  |  | 0
331 | SenNomVolt | Nominal sensor voltage | Nominal sensor voltage (used for sensor temperature voltage compensation) | 65 | 26 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
332 | VoltTCor | Sensor voltage compensation factor | Sensor voltage compensation factor (used for sensor temperature voltage comp.) | 116 | 16 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 9 |  |  | 0
333 | PCBSerial | PCB serial number | PCB serial number (unique number or string written by PCB manufacturer) | 119 | 31 | c |  |  |  |  | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 4 |  |  | 0
334 | MinMTime | Minimum measure time | Minimum measure time between sensors of a piston prover in 0.1 seconds | 115 | 24 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 20 |  |  | 0
335 | Bus1Parity | Fieldbus1 parity | Fieldbus1 parity (main interface) | 125 | 12 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 23 |  |  | 3 | 0: None ¦ 1: Odd ¦ 2: Even
336 | Bus2Parity | Fieldbus2 parity | Fieldbus2 parity | 124 | 12 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 23 |  |  | 3 | 0: None ¦ 1: Odd ¦ 2: Even
337 | FirmwId | Firmware id | Firmware identification number | 0 | 8 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 12 |  |  | 0
338 | Valve100 | Valve 100% | Valve 100% current/voltage | 114 | 13 | f |  | 0 | 0 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 8 |  |  | 0
339 | SetpMin | Setpoint minimum | Minimum setpoint for controller | 114 | 14 | f |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
340 | PCBSerialI | PCB serial number index | PCB serial number index | 119 | 30 | c |  |  | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 4 | 0 | 0 | 2 | 0: Main PCB serial number ¦ 1: Sensor PCB serial number
341 | Bus1MAC | Fieldbus1 MAC address | Fieldbus1 MAC address | 125 | 13 | c |  | 00-00-00-00-00-00 |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 23 |  |  | 0
342 | FluidNrSrc | Source fluidset index | Index of source fluidset for conversion | 126 | 0 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
343 | CapRanCalM | Capacity range calculation method | Calculation method for capacity range | 126 | 1 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
344 | ActFlPIdx | Actual fluid property index | Actual fluid property index | 126 | 2 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 4 | 0: Density ¦ 1: Heat capacity ¦ 2: Thermal conductivity ¦ 3: Viscosity
345 | ActFlPCalM | Actual fluid property calculation method | Actual fluid property calculation method | 126 | 3 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
346 | MixFrcType | Mix fraction type | Mix fraction type | 126 | 4 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 3 | 0: Volume fractions ¦ 1: Mass fractions ¦ 2: Mole fractions
347 | MixVolTemp | Mix volume temperature | Mix volume type temperature (�C) | 126 | 5 | f |  | 0 | -273.15 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
348 | MixVolPres | Mix volume pressure | Mix volume type pressure (bar (a)) | 126 | 6 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
349 | MixCompIdx | Mix component index | Mix component index | 126 | 7 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
350 | MixCompFrc | Mix component fraction | Mix component fraction | 126 | 8 | f |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
351 | MixCompNam | Mix component fluid name | Mix component fluid name | 126 | 9 | c |  | AIR |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 3 |  |  | 0
352 | EqMFldName | Equivalent measure fluid name | Fluid name for equivalent measure | 126 | 10 | c |  | AIR |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 19 |  |  | 0
353 | EqMeasure | Equivalent measure | Equivalent measure at actual conditions (kg/s) | 126 | 11 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 19 |  |  | 0
354 | FlConCondS | Fluid conversion conditions selection | Fluid conversion conditions (temperature and pressure) selection | 126 | 12 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 4 | 0: Temperature from fluidset temperature parameter (p181) ¦ 0: Pressure from fluidset pressure inlet parameter (p178) ¦ 1: Temperature from temperature sensor parameter (p142) ¦ 2: Pressure from pressure sensor parameter (p143)
355 | SensCapMin | Sensor+restriction capacity minimum | Sensor+restriction capacity minimum | 126 | 13 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
356 | SensCapMax | Sensor+restriction capacity maximum | Sensor+restriction capacity maximum | 126 | 14 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
357 | ValvCapMin | Valve capacity minimum | Valve capacity minimum | 126 | 15 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
358 | ValvCapMax | Valve capacity maximum | Valve capacity maximum | 126 | 16 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 3 |  |  | 0
359 | MaxDosTime | Maximum allowed dosing time | Maximum allowed dosing time (s) | 104 | 15 | f |  | 0 | 0 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
360 | DosingTime | Most recent dosing time | Most recent dosing time (s) | 104 | 16 | f |  | 0 | 0 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
361 | CtrlHyst | Controller hysteresis | Controller hysteresis | 114 | 15 | f |  | 0 | 0 | 1 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 8 |  |  | 0
362 | PressInZA | Pressure sensor input zero adjust | Pressure sensor input, zero adjust | 126 | 17 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
363 | PressInSA | Pressure sensor input span adjust | Pressure sensor input, span adjust | 126 | 18 | f |  | 1 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 10 |  |  | 0
364 | SpecialIdx | Special parameter index | Special parameter index | 127 | 0 | l |  | 0 | 0 | 2147483648 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
365 | SpecLong | Special parameter i long | Special parameter i long | 127 | 1 | l |  | 0 | -2147483648 | 2147483648 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
366 | SpecFloat | Special parameter i float | Special parameter i float | 127 | 2 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
367 | SpecName | Special parameter i name | Special parameter i name | 127 | 3 | c |  |  |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
368 | SpecLong1 | Special parameter l1 long | Special parameter l1 long | 127 | 4 | l |  | 0 | -2147483648 | 2147483648 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
369 | SpecLong2 | Special parameter l2 long | Special parameter l2 long | 127 | 5 | l |  | 0 | -2147483648 | 2147483648 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
370 | SpecLong3 | Special parameter l3 long | Special parameter l3 long | 127 | 6 | l |  | 0 | -2147483648 | 2147483648 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
371 | SpecLong4 | Special parameter l4 long | Special parameter l4 long | 127 | 7 | l |  | 0 | -2147483648 | 2147483648 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
372 | SpecLong5 | Special parameter l5 long | Special parameter l5 long | 127 | 8 | l |  | 0 | -2147483648 | 2147483648 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
373 | SpecFloat1 | Special parameter f1 float | Special parameter f1 float | 127 | 9 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
374 | SpecFloat2 | Special parameter f2 float | Special parameter f2 float | 127 | 10 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
375 | SpecFloat3 | Special parameter f3 float | Special parameter f3 float | 127 | 11 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
376 | SpecFloat4 | Special parameter f4 float | Special parameter f4 float | 127 | 12 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
377 | SpecFloat5 | Special parameter f5 float | Special parameter f5 float | 127 | 13 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
378 | BusIntIdx | Fieldbus interface index | Fieldbus interface index | 125 | 7 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 23 |  |  | 5 | 0: Fieldbus interface (main interface, fieldbus1) ¦ 1: Instrument interface (fieldbus2) ¦ 2: Display interface ¦ 3: Service interface ¦ 4: HART interface
379 | SelfStat | Selftest status | Selftest status | 119 | 7 | l |  | 0 | 0 | 2147483648 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 4 | 1: Production test completed succesfully ¦ 2: Service test completed ¦ 4: Production test detection enabled ¦ 8: Service test detection enabled
380 | SelfBIdx | Selftest block index | Selftest block index | 119 | 8 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
381 | SelfTIdx | Selftest test index | Selftest test index | 119 | 9 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
382 | SelfName | Selftest test description/name | Selftest test description/name | 119 | 10 | c |  | 0 |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
383 | SelfRes | Selftest test result | Selftest test result (pass/fail) | 119 | 11 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
384 | SelfLower | Selftest test lower limit | Selftest test lower limit | 119 | 12 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
385 | SelfUpper | Selftest test upper limit | Selftest test upper limit | 119 | 13 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
386 | SelfLow | Selftest test lowest value | Selftest test lowest value | 119 | 14 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
387 | SelfHigh | Selftest test highest value | Selftest test highest value | 119 | 15 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
388 | SelfAvg | Selftest test average value | Selftest test average value | 119 | 16 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
389 | SelfStDv | Selftest test standard deviation | Selftest test standard deviation | 119 | 17 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 |  |  | 0
390 | Bus1IP | Fieldbus1 IP address | IP address, when no address configured value is 0.0.0.0 (DHCP) | 125 | 14 | c |  | 0.0.0.0 |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 0
391 | Bus1SubMas | Fieldbus1 subnet mask | Subnet mask for the IP address | 125 | 15 | c |  | 0.0.0.0 |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 0
392 | Bus1GateW | Fieldbus1 gateway address | IP address of the gateway | 125 | 16 | c |  | 0.0.0.0 |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 |  |  | 0
393 | TotalzrVal | Totalizer value | Totalizer value | 104 | 17 | f |  | 0 | -3.4e+40 | 3.4e+40 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
394 | TotalzrUnt | Totalizer unit | Totalizer unit | 104 | 18 | c |  | ln/min |  |  | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 16 |  |  | 0
395 | CntrThresh | Counter threshold | Counter threshold | 104 | 20 | f |  | 0 | -3.4e+40 | 3.4e+40 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 |  |  | 0
396 | Bus1Ver | Fieldbus1 version | Contains fieldbus firmware version information | 125 | 22 | c |  |  |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 23 |  |  | 0
397 | RelHum | Relative humidity | Relative humidity (0..100%) | 65 | 20 | f |  | 0 | 0 | 100 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 9 |  |  | 0
398 | DosType | Dosing type | Dosing type | 112 | 1 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | 16 |  |  | 3 | 0: Continuous ¦ 1: Ratio ¦ 2: Batch
399 | DosCtrType | Dosing controller type | Dosing controller type | 112 | 2 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | 16 | 0 | 0 | 2 | 0: PID controlled ¦ 1: On/off controlled
400 | DosActType | Dosing rejection mode | Dosing rejection type | 112 | 3 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 2 | 0: Pump/control valve ¦ 1: On/off valve
401 | DosStartTp | Dosing mode | Dosing mode | 112 | 4 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 2 | 0: External trigger [dosing type=batch] ¦ 1: Time interval [dosing type=batch]
402 | DosDelayT | Dosing start delay time | Dosing start delay time | 112 | 5 | f |  | 0 | 0 | 3.4e+40 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
403 | DosTime | Dosing time | Dosing time | 112 | 6 | f |  | 0 | 0 | 3.4e+40 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
404 | DosRepInt | Dosing repetition interval | Dosing repetition interval | 112 | 7 | f |  | 0 | 0 | 3.4e+40 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
405 | DosSize | Dosing size | Dosing size | 112 | 8 | f |  | 0 | 0 | 3.4e+40 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
406 | DosAcc | Dosing accuracy [%] | Dosing accuracy [%] | 112 | 9 | f |  | 0 | 0 | 3.4e+40 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
407 | ActDosSze | Actual dosing size | Actual dosing size | 112 | 10 | f |  | 0 | 0 | 3.4e+40 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
408 | ActDosTime | Actual dosing time [s] | Actual dosing dosing time [s] | 112 | 11 | f |  | 0 | 0 | 3.4e+40 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
409 | ActDosAcc | Actual dosing accuracy [%] | Actual dosing accuracy [%] | 112 | 12 | f |  | 0 | -3.4e+40 | 3.4e+40 | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
410 | DosUnt | Dosing unit | Dosing unit | 112 | 0 | c |  | 0 |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 16 | 0 | 0 | 0
411 | DiaNEvtIdx | Diagnostic newest event index | Diagnostic newest event index | 118 | 14 | i |  | 0 | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
412 | DiaEvtIdx | Diagnostic event index | Diagnostic event index | 118 | 15 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
413 | DiaEvtCode | Diagnostic event code | Diagnostic event code | 118 | 16 | i |  | 0 | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
414 | DiaEvtDesc | Diagnostic event description | Diagnostic event description | 118 | 20 | c |  | 0 |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
415 | DiaEvtAct | Diagnostic event active | Diagnostic event active | 118 | 17 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
416 | DiaEvtNSt | Diagnostic event namur status | Diagnostic event namur status | 118 | 18 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
417 | DiaEvtTime | Diagnostic event timestamp | Diagnostic event timestamp | 118 | 21 | l |  | 0 | 0 | 4294967295 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
418 | InstrNSt | Instrument namur status | Instrument namur status | 118 | 0 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
419 | InstrTime | Instrument timestamp | Instrument timestamp | 118 | 13 | l |  | 0 | 0 | 4294967295 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
420 | OpHisIdx | Operational history parameter index | Operational history parameter index | 118 | 26 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
421 | OpHisName | Operational history parameter name | Operational history parameter name | 118 | 27 | c |  | 0 |  |  | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
422 | OpHisMin | Operational history parameter minimum value | Operational history parameter minimum value | 118 | 28 | f |  | 0 | 0 | 3.4e+40 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
423 | OpHisMax | Operational history parameter maximum value | Operational history parameter maximum value | 118 | 29 | f |  | 0 | 0 | 3.4e+40 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
424 | OpHisAvg | Operational history parameter average | Operational history parameter average | 118 | 30 | f |  | 0 | 0 | 3.4e+40 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
425 | OpHisStdDv | Operational history parameter standard deviation | Operational history parameter standard deviation | 118 | 31 | f |  | 0 | 0 | 3.4e+40 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
426 | Bus1StopB | Fieldbus stop bits | Fieldbus stop bits | 125 | 23 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 | 0 | 0 | 0
427 | Bus1ConMod | Fieldbus connection mode | Fieldbus connection mode | 125 | 24 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 | 0 | 0 | 0
428 | Bus1PssKey | Fieldbus passkey | Fieldbus passkey | 125 | 25 | l |  | 0 | 0 | 4294967295 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 23 | 0 | 0 | 0
429 | FeatUnlckC | Feature unlock code | Feature unlock code | 119 | 27 | c |  |  |  |  | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 | 0 | 0 | 0
430 | FeatSel | Feature selection | Feature selection | 119 | 28 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 | 0 | 0 | 0
431 | FeatState | Feature state | Feature state | 119 | 29 | c |  | 0 | 0 | 255 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | 12 | 0 | 0 | 0
432 | CtrlFunc | Control function | Control function | 115 | 10 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 8 | 0 | 0 | 0
433 | ActIdx | Actuator index | Actuator index | 114 | 0 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | 8 | 0 | 0 | 0
434 | DosStatus | Dosing status | Dosing status | 112 | 13 | i |  |  | 0 | 65535 | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 24 | 0 | 0 | 0
435 | BodyCode | Body code | Body code | 114 | 9 | i |  | 0 | 0 | 65535 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
436 | BodyRevCod | Body configuration code | Body configuration code | 114 | 10 | c |  | 0 | 0 | 255 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 13 |  |  | 0
437 | DosSeqNr | Batch sequence number | Batch sequence number | 112 | 14 | l |  | 0 | 0 | 4294967295 | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | 16 |  |  | 0
438 | LookDp | Lookup table delta p | Lookup table for linearisation delta p | 65 | 27 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0
439 | LookP | Lookup table pressure | Lookup table for linearisation pressure | 65 | 28 | f |  | 0 | -3.4028e+38 | 3.4028e+38 | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | 3 |  |  | 0

## value_options
parameter | name | value | description | filter | id
6 | arbitrage | 1 | temporary busmaster |  | 6
6 | arbitrage | 2 | always busmaster |  | 342
6 | arbitrage | 3 | automatic busmaster |  | 417
6 | arbitrage | 67 | auto busmaster and auto bus optimalization (fast token ring) |  | 418
6 | arbitrage | 83 | auto busmaster and auto bus optimalization with fail safe |  | 584
12 | cntrlmode | 0 | setpoint = BUS setpoint |  | 12
12 | cntrlmode | 1 | setpoint = analog input |  | 140
12 | cntrlmode | 2 | setpoint = master output(FLOW-BUS) * slave factor(FLOW-BUS) |  | 141
12 | cntrlmode | 3 | close valve |  | 142
12 | cntrlmode | 4 | controller idle (no reaction on changes in sensor signal) |  | 143
12 | cntrlmode | 5 | testmode enable (select subject with par 70) |  | 144
12 | cntrlmode | 6 | tuningmode enable (select subtject with par 79) |  | 145
12 | cntrlmode | 7 | setpoint = 100% |  | 146
12 | cntrlmode | 8 | purge valve (fully open) |  | 147
12 | cntrlmode | 9 | calibration mode enable (select subject with par 58) |  | 148
12 | cntrlmode | 10 | setpoint = master output(analog in) * slave factor(FLOW-BUS) |  | 149
12 | cntrlmode | 11 | setpoint = keyboard OR FLOW-BUS setpoint |  | 150
12 | cntrlmode | 12 | setpoint = 0% |  | 151
12 | cntrlmode | 13 | setpoint = master output(FLOW-BUS) * slave factor(analog in) |  | 377
12 | cntrlmode | 14 | (FPP) Range select mode |  | 446
12 | cntrlmode | 15 | (FPP) Manual start sensor select, automatic end sensor |  | 447
12 | cntrlmode | 16 | (FPP) Automatic start sensor select, manual end sensor |  | 448
12 | cntrlmode | 17 | (FPP) Automatic start and end sensor |  | 449
12 | cntrlmode | 18 | setpoint = RS232 setpoint |  | 450
12 | cntrlmode | 19 | RS232 broadcast mode |  | 501
12 | cntrlmode | 20 | valve stearing (valve = setpoint) |  | 502
12 | cntrlmode | 21 | analog valve stearing (valve = analog setpoint) |  | 503
12 | cntrlmode | 22 | valve safe state |  | 570
12 | cntrlmode | 23 | Use fluidset conditions for conversion (instead of p354) |  | 616
12 | cntrlmode | 24 | vent DEPx (when available) |  | 604
22 | sensortype | 0 | pressure (controller) |  | 22
22 | sensortype | 1 | liquid volume (controller) |  | 152
22 | sensortype | 2 | liquid/gas mass (controller) |  | 153
22 | sensortype | 3 | gas volume (controller) |  | 154
22 | sensortype | 4 | other sensor type (controller) |  | 156
22 | sensortype | 5 | temperature (controller) |  | 588
22 | sensortype | 6 | time (duration) (controller) |  | 590
22 | sensortype | 7 | frequency (controller) |  | 592
22 | sensortype | 8 | weight (controller) |  | 594
22 | sensortype | 9 | solutions (controller |  | 596
22 | sensortype | 10 | density (controller) |  | 620
22 | sensortype | 128 | pressure (sensor) |  | 157
22 | sensortype | 129 | liquid volume (sensor) |  | 158
22 | sensortype | 130 | liquid/gas mass (sensor) |  | 159
22 | sensortype | 131 | gas volume (sensor) |  | 160
22 | sensortype | 132 | other sensor type (sensor) |  | 161
22 | sensortype | 133 | temperature (sensor) |  | 589
22 | sensortype | 134 | time (durarion) (sensor) |  | 591
22 | sensortype | 135 | frequency (sensor) |  | 593
22 | sensortype | 136 | weight (sensor) |  | 595
22 | sensortype | 137 | solutions (sensor) |  | 597
22 | sensortype | 138 | density (sensor) |  | 621
28 | alarminfo | 0 | no error message in alarm error status register | &H01 | 28
28 | alarminfo | 0 | no warning message in alarm warning status register | &H02 | 378
28 | alarminfo | 0 | no minimum alarm message (measure>minimum limit) | &H04 | 379
28 | alarminfo | 0 | no maximum alarm message (measure<maximum limit) | &H08 | 380
28 | alarminfo | 0 | batch counter has not reached its limit | &H10 | 381
28 | alarminfo | 0 | response O.K. (setpoint-measure within limit) | &H20 | 382
28 | alarminfo | 0 | master output signal O.K. (or not used) | &H40 | 383
28 | alarminfo | 0 | hardware O.K. | &H80 | 384
28 | alarminfo | 1 | at least 1 error message in alarm error status register | &H01 | 385
28 | alarminfo | 1 | at least 1 warning message in alarm warning status register | &H02 | 386
28 | alarminfo | 1 | minimum alarm message for measured signal | &H04 | 387
28 | alarminfo | 1 | maximum alarm message for measured signal | &H08 | 388
28 | alarminfo | 1 | batch counter has reached its limit | &H10 | 389
28 | alarminfo | 1 | response alarm message: setpoint-measure is too high | &H20 | 390
28 | alarminfo | 1 | master output signal not received: check master instrument | &H40 | 391
28 | alarminfo | 1 | hardware error message: check your hardware | &H80 | 392
44 | opermodeTA | 0 | OFF |  | 44
44 | opermodeTA | 1 | A: MAX & RESP AUTO; T: UP TO LIMIT |  | 205
44 | opermodeTA | 2 | A: MIN & RESP AUTO; T: UP AND REPEAT |  | 206
44 | opermodeTA | 3 | A: MAX & RESP; T: DOWN FROM LIMIT |  | 207
44 | opermodeTA | 4 | A: MIN & RESP; T: DOWN AND REPEAT |  | 208
44 | opermodeTA | 5 | A: MAXIMUM ALARM; T: ALWAYS UP |  | 209
44 | opermodeTA | 6 | A: MINIMUM ALARM |  | 210
44 | opermodeTA | 7 | A: RESPONSE ALARM |  | 211
53 | AnalogMode | 0 | 0...5 Vdc operation | &H3F | 53
53 | AnalogMode | 0 | Analog input enabled | &H40 | 460
53 | AnalogMode | 0 | Analog output enabled | &H80 | 462
53 | AnalogMode | 1 | 0...10 Vdc operation | &H3F | 218
53 | AnalogMode | 1 | Analog input disabled | &H40 | 461
53 | AnalogMode | 1 | Analog output disabled | &H80 | 463
53 | AnalogMode | 2 | 0...20 mA operation | &H3F | 219
53 | AnalogMode | 3 | 4...20 mA operation | &H3F | 220
53 | AnalogMode | 4 | 15...20 mA operation | &H3F | 221
58 | CalMode | 0 | idle: no action |  | 58
58 | CalMode | 1 | adc self calibration |  | 222
58 | CalMode | 2 | dmfc |  | 223
58 | CalMode | 3 | dmfc |  | 224
58 | CalMode | 4 | dmfc |  | 225
58 | CalMode | 5 | dmfc |  | 226
58 | CalMode | 6 | dmfc |  | 227
58 | CalMode | 7 | dmfc |  | 228
58 | CalMode | 8 | dmfc |  | 229
58 | CalMode | 9 | customer zero |  | 230
58 | CalMode | 10 | adjust Vref output by connecting it to analog in |  | 231
58 | CalMode | 11 | adjust analog out by connecting it to analog in |  | 232
58 | CalMode | 12 | adjust valveoutput by connecting it to analog in |  | 233
58 | CalMode | 13 | dmfc |  | 234
58 | CalMode | 14 | dmfc |  | 235
58 | CalMode | 15 | analog output = 0 % |  | 162
58 | CalMode | 16 | analog output = 100 % |  | 163
58 | CalMode | 17 | analog output = 50 % |  | 269
58 | CalMode | 18 | factory zero |  | 494
58 | CalMode | 19 | sensor differentiator (setpoint steps are needed!) |  | 516
58 | CalMode | 20 | automatic sensor configuration |  | 567
58 | CalMode | 21 | sensor temperature calibration |  | 568
58 | CalMode | 22 | customer zero (no control mode 9 needed) |  | 569
58 | CalMode | 255 | Error mode (result of previous cal mode) |  | 517
60 | monitor | 0 | (filtered) setpoint |  | 60
60 | monitor | 1 | controller error input signal / raw sensor signal |  | 167
60 | monitor | 2 | controller output signal to valve |  | 168
60 | monitor | 3 | sensor signal slow |  | 169
60 | monitor | 4 | sensor signal slow filtered |  | 170
60 | monitor | 5 | linearization output |  | 171
60 | monitor | 6 | differentiator output |  | 172
60 | monitor | 7 | differentiator output filtered |  | 173
60 | monitor | 8 | normal sensor signal (Output) |  | 164
60 | monitor | 9 | analog input signal |  | 165
60 | monitor | 10 | power supply voltage |  | 166
60 | monitor | 11 | mass flow in display unit (normally ln/min) |  | 271
60 | monitor | 12 | volume flow in l/min |  | 272
60 | monitor | 13 | temperature in �C |  | 273
60 | monitor | 14 | pressure absolute in mbara |  | 274
60 | monitor | 15 | time in msec/frequency in Hz. |  | 275
60 | monitor | 16 | calibrated volume at actual sensor in ml |  | 276
60 | monitor | 17 | delta-P pressure in mbarg |  | 277
60 | monitor | 18 | atmospheric (barometer) pressure in mbara |  | 278
60 | monitor | 19 | mass flow in kg/min |  | 458
61 | AlarmReg1 | 0 | No diagnostics available in warning register | &H8000000000000000 | 466
61 | AlarmReg1 | 1 | Diagnostics available in warning register | &H8000000000000000 | 464
62 | AlarmReg2 | 0 | No diagnostics available in error register | &H8000000000000000 | 470
62 | AlarmReg2 | 1 | Diagnostics available in error register | &H8000000000000000 | 468
67 | ADCcntrReg | 0 | ADC gain = 1x | &H1C0000 | 356
67 | ADCcntrReg | 0 | ADC bipolar mode | &H001000 | 374
67 | ADCcntrReg | 0 | Disable zero measure threshold | &H1000000 | 520
67 | ADCcntrReg | 1 | ADC gain = 2x | &H1C0000 | 357
67 | ADCcntrReg | 1 | ADC unipolar mode | &H001000 | 375
67 | ADCcntrReg | 1 | Enable zero measure threshold | &H1000000 | 519
67 | ADCcntrReg | 2 | ADC gain = 4x | &H1C0000 | 358
67 | ADCcntrReg | 3 | ADC gain = 8x | &H1C0000 | 359
67 | ADCcntrReg | 4 | ADC gain = 16x | &H1C0000 | 360
67 | ADCcntrReg | 5 | ADC gain = 32x | &H1C0000 | 361
67 | ADCcntrReg | 6 | ADC gain = 64x | &H1C0000 | 362
67 | ADCcntrReg | 7 | ADC gain = 128x | &H1C0000 | 363
69 | AlarmEnble | 0 | disable |  | 69
69 | AlarmEnble | 1 | enable |  | 343
70 | TestMode | 0 | idle; no action |  | 70
70 | TestMode | 1 | uProcessor |  | 236
70 | TestMode | 2 | IO |  | 237
70 | TestMode | 3 | RAM |  | 238
70 | TestMode | 4 | FRAM |  | 239
70 | TestMode | 5 | ADC |  | 240
70 | TestMode | 6 | DAC |  | 241
70 | TestMode | 7 | sensor |  | 242
70 | TestMode | 8 | valve drive circuit |  | 243
70 | TestMode | 9 | Vref |  | 244
70 | TestMode | 10 | FLOW-BUS |  | 245
70 | TestMode | 11 | calibration |  | 246
70 | TestMode | 12 | keyboard |  | 247
71 | ChanSelect | 1 | AD channel 1 |  | 71
71 | ChanSelect | 2 | AD channel 2 |  | 344
79 | TuningMode | 0 | idle; no action |  | 79
79 | TuningMode | 1 | sensor |  | 248
79 | TuningMode | 2 | valve |  | 249
79 | TuningMode | 3 | Fuzzy controller normal operation |  | 250
79 | TuningMode | 4 | Fuzzy controller open at zero |  | 251
79 | TuningMode | 5 | PID controller |  | 410
80 | DefVlvType | 0 | normally closed |  | 80
80 | DefVlvType | 1 | normally opened |  | 252
80 | DefVlvType | 2 | normally closed inverse controlled |  | 266
80 | DefVlvType | 3 | normally opened inverse controlled |  | 267
80 | DefVlvType | 4 | remain position |  | 253
86 | IOStatus | 1 | read diagnostic jumper (no diagnostics, read/write) | &H01 | 254
86 | IOStatus | 1 | not used | &H02 | 255
86 | IOStatus | 1 | read micro switch (read/write) | &H08 | 257
86 | IOStatus | 1 | diagnostic jumper set (read only) | &H10 | 260
86 | IOStatus | 1 | initialization jumper set (read only) | &H20 | 261
86 | IOStatus | 1 | analog jumper set (read only) | &H40 | 262
86 | IOStatus | 1 | micro switch pressed (read only) | &H80 | 263
86 | IOStatus | 1 | read analog jumper (use cntrlmode, read/write) | &H04 | 522
106 | PressSensr | 0 | delta-P 0..5" W.C. |  | 106
106 | PressSensr | 1 | delta-P 0...10" W.C. |  | 264
106 | PressSensr | 2 | absolute pressure 800-1200 mbar |  | 265
106 | PressSensr | 3 | absolute pressure 800-1100 mbar |  | 270
106 | PressSensr | 4 | delta-P -5...0 "W.C. |  | 399
106 | PressSensr | 5 | delta-P -10...0 "W.C. |  | 400
106 | PressSensr | 6 | delta-P -10...+10 "W.C. |  | 401
106 | PressSensr | 7 | delta-P 0...1 PSI |  | 402
106 | PressSensr | 8 | delta-P -1...0 PSI |  | 403
106 | PressSensr | 9 | absolute pressure 0-10 bar |  | 527
114 | Reset | 0 | no reset |  | 114
114 | Reset | 1 | reset counter value (no mode change) or common reset |  | 174
114 | Reset | 2 | reset alarm |  | 175
114 | Reset | 3 | restart batch counter |  | 177
114 | Reset | 4 | reset counter value (counter off) |  | 176
114 | Reset | 5 | Reset module (soft reset) |  | 413
118 | AlrmMode | 0 | off |  | 118
118 | AlrmMode | 1 | alarm on absolute limits |  | 178
118 | AlrmMode | 2 | alarm on limits related to setpoint (response alarm) |  | 179
118 | AlrmMode | 3 | alarm when instrument powers-up (eg. after power-down) |  | 416
119 | AlrmOutMod | 0 | no relais activity at alarm |  | 119
119 | AlrmOutMod | 1 | relais pulses until reset |  | 180
119 | AlrmOutMod | 2 | relais activated until reset |  | 181
120 | AlrmStpMod | 0 | no setpoint change at alarm |  | 120
120 | AlrmStpMod | 1 | new/safe setpoint at alarm enabled (set at par 121) |  | 182
125 | CntrOutMod | 0 | no relais activity at batch limit |  | 125
125 | CntrOutMod | 1 | relais pulses after reaching batch limit until reset |  | 183
125 | CntrOutMod | 2 | relais activated after reaching batch limit until reset |  | 184
126 | CntrStpMod | 0 | setpoint change at batch limit disabled |  | 126
126 | CntrStpMod | 1 | setpoint change at batch limit enabled |  | 493
130 | CntrMode | 0 | off |  | 130
130 | CntrMode | 1 | counting upwards continuously |  | 185
130 | CntrMode | 2 | counting up to limit (batchcounter) |  | 186
147 | rangeselct | 0 | calibration ready/stop |  | 279
147 | rangeselct | 1 | run calibration until stopsensor 1/select range 1 |  | 280
147 | rangeselct | 2 | run calibration until stopsensor 2/select range 2 |  | 281
147 | rangeselct | 3 | run calibration until stopsensor 3/select range 3 |  | 282
147 | rangeselct | 4 | run calibration until stopsensor 4/select range 4 |  | 283
147 | rangeselct | 5 | run calibration and select range 5 |  | 284
147 | rangeselct | 9 | run calibration with automatic range selection |  | 285
147 | rangeselct | 19 | run until stopsensor 1 until 3 values between limit |  | 286
147 | rangeselct | 29 | run until stopsensor 2 until 3 values between limit |  | 287
147 | rangeselct | 39 | run until stopsensor 3 until 3 values between limit |  | 288
147 | rangeselct | 49 | run until stopsensor 4 until 3 values between limit |  | 289
147 | rangeselct | 59 | run and select range 5 until 3 values between limit |  | 290
147 | rangeselct | 99 | run with auto-select + 3 values between limit |  | 291
156 | RstAlarmEn | 0 | no reset possible |  | 300
156 | RstAlarmEn | 1 | reset: keyboard |  | 301
156 | RstAlarmEn | 2 | reset: external |  | 302
156 | RstAlarmEn | 3 | reset: keyboard or external |  | 303
156 | RstAlarmEn | 4 | reset: FLOW-BUS |  | 304
156 | RstAlarmEn | 5 | reset: FLOW-BUS or keyboard |  | 305
156 | RstAlarmEn | 6 | reset: FLOW-BUS or external |  | 306
156 | RstAlarmEn | 7 | reset: FLOW-BUS or keyboard or external |  | 307
156 | RstAlarmEn | 8 | reset: automatic |  | 308
156 | RstAlarmEn | 9 | reset: automatic or keyboard |  | 309
156 | RstAlarmEn | 10 | reset: automatic or external |  | 310
156 | RstAlarmEn | 11 | reset: automatic or keyboard or external |  | 311
156 | RstAlarmEn | 12 | reset: automatic or FLOW-BUS |  | 312
156 | RstAlarmEn | 13 | reset: automatic or FLOW-BUS or keyboard |  | 313
156 | RstAlarmEn | 14 | reset: automatic or FLOW-BUS or external |  | 314
156 | RstAlarmEn | 15 | reset: automatic or FLOW-BUS or keyboard or external |  | 315
157 | RstCountEn | 0 | no reset possible |  | 334
157 | RstCountEn | 1 | reset: keyboard |  | 335
157 | RstCountEn | 2 | reset: external |  | 336
157 | RstCountEn | 3 | reset: keyboard or external |  | 337
157 | RstCountEn | 4 | reset: FLOW-BUS |  | 338
157 | RstCountEn | 5 | reset: FLOW-BUS or keyboard |  | 339
157 | RstCountEn | 6 | reset: FLOW-BUS or external |  | 340
157 | RstCountEn | 7 | reset: FLOW-BUS or keyboard or external |  | 341
157 | RstCountEn | 8 | reset: automatic |  | 419
157 | RstCountEn | 9 | reset: automatic or keyboard |  | 420
157 | RstCountEn | 10 | reset: automatic or external |  | 421
157 | RstCountEn | 11 | reset: automatic or keyboard or external |  | 422
157 | RstCountEn | 12 | reset: automatic or FLOW-BUS |  | 423
157 | RstCountEn | 13 | reset: automatic or FLOW-BUS or keyboard |  | 424
157 | RstCountEn | 14 | reset: automatic or FLOW-BUS or external |  | 425
157 | RstCountEn | 15 | reset: automatic or FLOW-BUS or keyboard or external |  | 426
166 | ContrType | 0 | valve in normal position after startup | &H01 | 393
166 | ContrType | 0 | open from zero with PID output to valve | &H02 | 395
166 | ContrType | 0 | fixed monitor output signal | &H04 | 397
166 | ContrType | 0 | voltage drift compensation for valve output turned on | &H08 | 414
166 | ContrType | 0 | auto slope disabled | &H10 | 427
166 | ContrType | 0 | automatic correction for valve open turned on | &H20 | 451
166 | ContrType | 0 | controller special mode (valve output steps) turned off | &H40 | 453
166 | ContrType | 0 | valve overshoot protection turned off | &H80 | 524
166 | ContrType | 1 | valve in safe position after startup | &H01 | 394
166 | ContrType | 1 | open from zero with ramp output to valve | &H02 | 396
166 | ContrType | 1 | monitor output changed at setpoint steps | &H04 | 398
166 | ContrType | 1 | voltage drift compensation for valve output turned off | &H08 | 415
166 | ContrType | 1 | auto slope enabled for pilot valves | &H10 | 428
166 | ContrType | 1 | automatic correction for valve open turned off | &H20 | 452
166 | ContrType | 1 | controller special mode (valve output steps) turned on | &H40 | 454
166 | ContrType | 1 | valve overshoot protection turned on | &H80 | 523
175 | IdenrNr | 0 | UFO?: Unidentified FLOW-BUS Object |  | 521
175 | IdentNr | 1 | RS232/FLOW-BUS interface |  | 405
175 | IdentNr | 2 | PC(ISA) interface |  | 187
175 | IdentNr | 3 | ADDA4 (4 channels) |  | 188
175 | IdentNr | 4 | R/C-module, 32 channels |  | 189
175 | IdentNr | 5 | T/A-module |  | 190
175 | IdentNr | 6 | ADDA1: 1 channel ADDA converter module |  | 191
175 | IdentNr | 7 | DMFC: digital mass flow controller |  | 192
175 | IdentNr | 8 | DMFM: digital mass flow meter |  | 193
175 | IdentNr | 9 | DEPC: digital electronic pressure controller |  | 194
175 | IdentNr | 10 | DEPM: digital electronic pressure meter |  | 203
175 | IdentNr | 11 | ACT: single actuator |  | 195
175 | IdentNr | 12 | DLFC: digital liquid flow controller |  | 196
175 | IdentNr | 13 | DLFM: digital liquid flow meter |  | 197
175 | IdentNr | 14 | DSCM-A: digital single channel module for analog instruments |  | 198
175 | IdentNr | 15 | DSCM-D: digital single channel module for digital instr. |  | 199
175 | IdentNr | 16 | FRM: FLOW-BUS rotor meter (calibration-instrument) |  | 200
175 | IdentNr | 17 | FTM: FLOW-BUS turbine meter (calibration-instrument) |  | 201
175 | IdentNr | 18 | FPP: FLOW-BUS piston prover/tube (calibration-instrument) |  | 202
175 | IdentNr | 19 | F/A-module: special version of T/A-module |  | 204
175 | IdentNr | 20 | DSCM-E: evaporator controller module (single channel) |  | 406
175 | IdentNr | 21 | DSCM-C: digital single channel module for calibrators |  | 407
175 | IdentNr | 22 | DDCM-A: digital dual channel module for analog instruments |  | 408
175 | IdentNr | 23 | DMCM-D: digital multi channel module for digital instruments |  | 409
175 | IdentNr | 24 | Profibus-DP/FLOW-BUS interface module |  | 455
175 | IdentNr | 25 | FLOW-BUS Coriolis Meter |  | 456
175 | IdentNr | 26 | FBI: FLOW-BUS Balance Interface |  | 457
175 | IdentNr | 27 | CORIFC: CoriFlow Controller |  | 490
175 | IdentNr | 28 | CORIFM: CoriFlow Meter |  | 491
175 | IdentNr | 29 | FICC: FLOW-BUS Interface Climate Control |  | 497
175 | IdentNr | 30 | IFI: Instrument FLOW-BUS Interface |  | 498
175 | IdentNr | 31 | KFI: Keithley FLOW-BUS Interface |  | 499
175 | IdentNr | 32 | FSI: FLOW-BUS Switch Interface |  | 500
175 | IdentNr | 33 | MSCI: Multi-Sensor/Confroller Interface |  | 518
175 | IdentNr | 34 | APP-D: Active Piston Prover |  | 525
175 | IdentNr | 35 | LFI: Leaktester FLOW-BUS Interface |  | 526
175 | IdentNr | 36 | Digital Batch Flow Controller |  | 602
175 | IdentNr | 37 | Digital PID Controller |  | 603
175 | IdentNr | 38 | Bus Gateway |  | 617
175 | IdentNr | 39 | DTC: digital temperature controller |  | 618
175 | IdentNr | 40 | DTM: digital temperature meter |  | 619
175 | IdentNr | 41 | DC: digital controller |  | 627
175 | IdentNr | 42 | DM: digital meter |  | 628
185 | DeviceFunc | 0 | Unknown |  | 429
185 | DeviceFunc | 1 | Interface |  | 430
185 | DeviceFunc | 2 | ADDA |  | 431
185 | DeviceFunc | 3 | Operator |  | 432
185 | DeviceFunc | 4 | Supervisor (totalizer/alarm) |  | 433
185 | DeviceFunc | 5 | Controller |  | 434
185 | DeviceFunc | 6 | Meter |  | 435
185 | DeviceFunc | 7 | Special |  | 436
185 | DeviceFunc | 8 | (Protocol) converter |  | 437
197 | CalType | 0 | Automatic capacity setting for optimal resolution | &H01 | 438
197 | CalType | 0 | Barometer value input via parameter 107: BaroPress | &H02 | 440
197 | CalType | 1 | Manual capacity setting for optimal resolution | &H01 | 439
197 | CalType | 1 | Barometer is master; input automatically from master | &H02 | 441
200 | InterfConf | 0 | Configuration A: 14 ch. Standard parms. with network scan |  | 442
200 | InterfConf | 1 | Configuration B: 14 ch. Standard parms with fixed chan list |  | 443
200 | InterfConf | 2 | Configuration C: 7 ch. Extended parms with fixed chan list |  | 444
200 | InterfConf | 3 | Configuration D: 11 ch. Extended parms with network scan |  | 459
208 | Mstatus | 0 | No diagnostics available in manufacturer status register | &H800000 | 484
208 | Mstatus | 1 | Diagnostics available in manufacturer status register | &H800000 | 485
209 | Mwarning | 0 | No diagnostics available in manufacturer warning register | &H800000 | 480
209 | Mwarning | 1 | Diagnostics available in manufacturer warning register | &H800000 | 481
210 | Merror | 0 | No diagnostics available in manufacturer error register | &H800000 | 488
210 | Merror | 1 | Diagnostics available in manufacturer error register | &H800000 | 489
212 | DiagMode | 0 | Debug mode off |  | 472
212 | DiagMode | 1 | Debug mode on |  | 473
213 | MStatEnabl | 0 | set status bit (range 0�127) |  | 474
213 | MStatEnabl | 127 | set status bit (range 0�127) |  | 475
213 | MStatEnabl | 254 | clear all status bits |  | 476
213 | MStatEnabl | 255 | set all status bits |  | 477
232 | ValveMode | 0 | voltage drive mode |  | 495
232 | ValveMode | 1 | current drive mode |  | 496
232 | ValveMode | 2 | 0V (for PIO use) |  | 583
238 | FldSetProp | 0 | Fluidset is disabled | &H01 | 507
238 | FldSetProp | 0 | Fluidset is not factory set | &H02 | 508
238 | FldSetProp | 0 | Fluidset is not calibrated on actual gas | &H04 | 509
238 | FldSetProp | 1 | Fluidset is enabled | &H01 | 504
238 | FldSetProp | 2 | Fluidset is factory set | &H02 | 505
238 | FldSetProp | 4 | Fluidset is calibrated on actual gas | &H04 | 506
295 | SensBridge | 1 | Bridge on | &H01 | 528
295 | SensBridge | 1 | 3 windings C | &H02 | 529
295 | SensBridge | 1 | 3 windings D | &H04 | 530
295 | SensBridge | 1 | Sensor protection enabled | &H200 | 531
295 | SensBridge | 1 | Automatic sensor configuration on | &H100 | 540
301 | ValveSafe | 0 | 0 mA |  | 533
301 | ValveSafe | 1 | max mA |  | 534
301 | ValveSafe | 2 | Close |  | 535
301 | ValveSafe | 3 | Open |  | 536
301 | ValveSafe | 4 | Idle |  | 537
301 | ValveSafe | 5 | Value (for DeviceNet only) |  | 538
305 | Bus1Select | 0 | FLOW-BUS |  | 541
305 | Bus1Select | 1 | Modbus |  | 542
305 | Bus1Select | 2 | ProPar |  | 543
305 | Bus1Select | 10 | DeviceNet |  | 572
305 | Bus1Select | 11 | EtherCAT |  | 606
305 | Bus1Select | 13 | PROFIBUS DP |  | 571
305 | Bus1Select | 14 | PROFINET |  | 605
305 | Bus1Select | 255 | NoBus |  | 573
306 | Bus1Medium | 0 | RS232 |  | 544
306 | Bus1Medium | 1 | RS485 |  | 545
307 | Bus2Mode | 0 | Normal |  | 546
307 | Bus2Mode | 1 | Config mode |  | 547
308 | Bus2Select | 0 | FLOW-BUS |  | 549
308 | Bus2Select | 1 | Modbus |  | 550
308 | Bus2Select | 2 | ProPar |  | 551
311 | Bus2Medium | 0 | RS232 |  | 553
311 | Bus2Medium | 1 | RS485 |  | 552
314 | PIOChSel | 0 | Analog input |  | 554
314 | PIOChSel | 1 | Analog output |  | 555
314 | PIOChSel | 2 | General purpose in-/output |  | 556
319 | PIOCfgSel | 0 | Voltage output (0..10 V) |  | 557
319 | PIOCfgSel | 1 | Current output (0..20 mA) |  | 558
319 | PIOCfgSel | 2 | Digital output |  | 559
319 | PIOCfgSel | 3 | Frequency output |  | 560
319 | PIOCfgSel | 4 | Duty cycle output (20 kHz) |  | 561
319 | PIOCfgSel | 5 | Digital pulse output |  | 562
319 | PIOCfgSel | 6 | Voltage input (0..10 V) |  | 563
319 | PIOCfgSel | 7 | Current input (0..20 mA) |  | 564
319 | PIOCfgSel | 8 | Digital input |  | 565
319 | PIOCfgSel | 255 | Disabled |  | 566
329 | SetMonitor | 0 | Setpoint |  | 574
329 | SetMonitor | 1 | Filtered setpoint |  | 575
329 | SetMonitor | 2 | Setpoint after linear slope |  | 576
335 | Bus1Parity | 0 | None |  | 577
335 | Bus1Parity | 1 | Odd |  | 578
335 | Bus1Parity | 2 | Even |  | 579
336 | Bus2Parity | 0 | None |  | 580
336 | Bus2Parity | 1 | Odd |  | 581
336 | Bus2Parity | 2 | Even |  | 582
340 | PCBSerialI | 0 | Main PCB serial number |  | 585
340 | PCBSerialI | 1 | Sensor PCB serial number |  | 586
344 | ActFlPIdx | 0 | Density |  | 611
344 | ActFlPIdx | 1 | Heat capacity |  | 612
344 | ActFlPIdx | 2 | Thermal conductivity |  | 613
344 | ActFlPIdx | 3 | Viscosity |  | 614
346 | MixFrcType | 0 | Volume fractions |  | 598
346 | MixFrcType | 1 | Mass fractions |  | 599
346 | MixFrcType | 2 | Mole fractions |  | 600
354 | FlConCondS | 0 | Temperature from fluidset temperature parameter (p181) | &H01 | 607
354 | FlConCondS | 0 | Pressure from fluidset pressure inlet parameter (p178) | &H02 | 609
354 | FlConCondS | 1 | Temperature from temperature sensor parameter (p142) | &H01 | 608
354 | FlConCondS | 2 | Pressure from pressure sensor parameter (p143) | &H02 | 610
378 | BusIntIdx | 0 | Fieldbus interface (main interface, fieldbus1) |  | 622
378 | BusIntIdx | 1 | Instrument interface (fieldbus2) |  | 623
378 | BusIntIdx | 2 | Display interface |  | 624
378 | BusIntIdx | 3 | Service interface |  | 625
378 | BusIntIdx | 4 | HART interface |  | 626
379 | SelfStat | 1 | Production test completed succesfully | &H01 | 629
379 | SelfStat | 2 | Service test completed | &H02 | 630
379 | SelfStat | 4 | Production test detection enabled | &H04 | 631
379 | SelfStat | 8 | Service test detection enabled | &H08 | 632
398 | DosType | 0 | Continuous |  | 633
398 | DosType | 1 | Ratio |  | 634
398 | DosType | 2 | Batch |  | 635
399 | DosCtrType | 0 | PID controlled |  | 636
399 | DosCtrType | 1 | On/off controlled |  | 637
400 | DosActType | 0 | Pump/control valve |  | 638
400 | DosActType | 1 | On/off valve |  | 639
401 | DosStartTp | 0 | External trigger [dosing type=batch] |  | 640
401 | DosStartTp | 1 | Time interval [dosing type=batch] |  | 641